## Cal(2018)
Ta 5:5 SELF CALIBRATION AT 53 Hz
Index | Date | Conditions, comments | Shunt | X | Y | Sh V | X stdev | Y stdev | SR830 | SR830 | [6 empty] | SD ppm | SD ppm | SD ppm | noise sd | noise sd
|  |  | Volts | V | V | stdev/100 | V | V | Gain | Reserve | Shunt | Com R | % load | X ppm | Y ppm |  | X | Y | shunt V | X | Y
1 | 2018-07-31 | 53 Hz zero burden on aux sec. (10 ohm common resistor). | 0.1001 | -2.9815e-05 | -7.5842e-06 | 1.7123e-07 | 4.6089e-08 | 4.3716e-08 | 13 | 2 | 0.2 | 10 | 10.0096 | -5.9574 | -1.5154 |  | 0.009209 | 0.0087348 | 0.00017107 | 0.019981 | 0.019981 | Initial Setup
2 |  |  | 0.1002 | 0.0012862 | 3.4279e-05 | 1.1172e-07 | 6e-08 | 1.1603e-07 | 18 | 2 | 0.2 | 10 | 10.0202 | 256.7263 | 6.842 |  | 0.011976 | 0.023159 | 0.0001115 | 0.01996 | 0.01996 | 1k ohm across main secondary
3 |  |  | 0.1003 | -4.2452e-05 | 0.00043209 | 1.05e-07 | 3.4592e-08 | 3.0594e-08 | 16 | 2 | 0.2 | 10 | 10.0302 | -8.4648 | 86.1583 |  | 0.0068976 | 0.0061004 | 0.00010468 | 0.01994 | 0.01994 | 1uF across main secondary
4 |  |  | 1.2509 | -0.00015484 | -0.00011511 | 3.9257e-06 | 2.355e-08 | 2.1967e-08 | 15 | 2 | 0.2 | 10 | 125.088 | -2.4758 | -1.8405 |  | 0.00037653 | 0.00035123 | 0.00031384 | 0.0015989 | 0.0015989 | Start at 125%. 3458A on 12V range.
5 |  |  | 1.199 | -0.00015283 | -0.00011127 | 2.1566e-06 | 4.3287e-08 | 3.4139e-08 | 15 | 2 | 0.2 | 10 | 119.9044 | -2.5492 | -1.856 |  | 0.00072202 | 0.00056944 | 0.00017986 | 0.001668 | 0.001668 | 120%. 3458A at 1.2V range
6 |  |  | 1.0007 | -0.00014166 | -9.4187e-05 | 1.879e-07 | 5.1223e-08 | 3.058e-08 | 15 | 2 | 0.2 | 10 | 100.0717 | -2.8311 | -1.8824 |  | 0.0010237 | 0.00061116 | 1.8777e-05 | 0.0019986 | 0.0019986
7 |  |  | 0.60239 | -0.00011465 | -5.2214e-05 | 5.4476e-07 | 5.0611e-08 | 4.4289e-08 | 15 | 2 | 0.2 | 10 | 60.2395 | -3.8066 | -1.7335 |  | 0.0016803 | 0.0014704 | 9.0433e-05 | 0.0033201 | 0.0033201
8 |  |  | 0.40046 | -9.0008e-05 | -3.2823e-05 | 3.7118e-07 | 7.504e-08 | 6.2759e-08 | 14 | 2 | 0.2 | 10 | 40.046 | -4.4952 | -1.6393 |  | 0.0037477 | 0.0031344 | 9.269e-05 | 0.0049943 | 0.0049943
9 |  |  | 0.20012 | -5.383e-05 | -1.5764e-05 | 2.5102e-07 | 5.1549e-08 | 4.8776e-08 | 14 | 2 | 0.2 | 10 | 20.0123 | -5.3797 | -1.5754 |  | 0.0051517 | 0.0048746 | 0.00012543 | 0.0099939 | 0.0099939
10 |  |  | 0.10004 | -3.002e-05 | -7.5752e-06 | 9.4637e-08 | 9.1432e-08 | 6.8576e-08 | 13 | 2 | 0.2 | 10 | 10.0041 | -6.0015 | -1.5144 |  | 0.018279 | 0.01371 | 9.4598e-05 | 0.019992 | 0.019992
11 |  |  | 0.050017 | -1.6063e-05 | -3.7569e-06 | 1.6022e-08 | 1.4992e-08 | 2.7332e-08 | 12 | 2 | 0.2 | 10 | 5.0017 | -6.4231 | -1.5022 |  | 0.0059949 | 0.010929 | 3.2033e-05 | 0.039986 | 0.039986
12 |  |  | 0.010301 | -3.5691e-06 | -7.8145e-07 | 8.3903e-08 | 3.2984e-08 | 4.5002e-08 | 10 | 2 | 0.2 | 10 | 1.0301 | -6.9295 | -1.5172 |  | 0.064039 | 0.087373 | 0.0008145 | 0.19415 | 0.19415 | Do a 1% point. 0.2 ohm in series with the Zetka.
13
14 | 2018-07-31 | 47 Hz zero burden on aux sec. (10 ohm common resistor). | 1.2516 | -0.00018194 | -0.00014881 | 3.3493e-06 | 4.0046e-08 | 4.165e-08 | 16 | 2 | 0.2 | 10 | 125.1583 | -2.9073 | -2.3779 |  | 0.00063993 | 0.00066556 | 0.0002676 | 0.001598 | 0.001598 | Start at 125%. 3458A on 12V range.
15 |  |  | 1.1994 | -0.00017936 | -0.00014391 | 1.909e-06 | 4.9714e-08 | 1.5667e-08 | 16 | 2 | 0.2 | 10 | 119.9406 | -2.9908 | -2.3997 |  | 0.00082897 | 0.00026125 | 0.00015916 | 0.0016675 | 0.0016675 | 120%. 3458A at 1.2V range
16 |  |  | 1.0001 | -0.00016626 | -0.00012256 | 1.8939e-07 | 5.0604e-08 | 4.0006e-08 | 15 | 2 | 0.2 | 10 | 100.0065 | -3.3249 | -2.4511 |  | 0.001012 | 0.00080006 | 1.8938e-05 | 0.0019999 | 0.0019999
17 |  |  | 0.60083 | -0.00013649 | -7.1187e-05 | 6.8711e-07 | 4.0652e-08 | 3.4752e-08 | 15 | 2 | 0.2 | 10 | 60.0832 | -4.5433 | -2.3696 |  | 0.0013532 | 0.0011568 | 0.00011436 | 0.0033287 | 0.0033287
18 |  |  | 0.40044 | -0.00010959 | -4.4985e-05 | 3.5264e-07 | 2.763e-08 | 3.5403e-08 | 15 | 2 | 0.2 | 10 | 40.0436 | -5.4733 | -2.2468 |  | 0.00138 | 0.0017682 | 8.8063e-05 | 0.0049946 | 0.0049946
19 |  |  | 0.20058 | -6.6739e-05 | -2.2069e-05 | 1.8009e-07 | 3.823e-08 | 2.8819e-08 | 14 | 2 | 0.2 | 10 | 20.058 | -6.6546 | -2.2006 |  | 0.0038119 | 0.0028735 | 8.9782e-05 | 0.0099711 | 0.0099711
20 |  |  | 0.10073 | -3.7682e-05 | -1.0956e-05 | 7.215e-08 | 5.4066e-08 | 3.532e-08 | 13 | 1 | 0.2 | 10 | 10.0727 | -7.4821 | -2.1753 |  | 0.010735 | 0.007013 | 7.163e-05 | 0.019856 | 0.019856
21 |  |  | 0.050832 | -2.0416e-05 | -5.4974e-06 | 3.3239e-08 | 3.7848e-08 | 2.9699e-08 | 12 | 1 | 0.2 | 10 | 5.0832 | -8.0326 | -2.163 |  | 0.014891 | 0.011685 | 6.5389e-05 | 0.039345 | 0.039345
22 |  |  | 0.010669 | -4.627e-06 | -1.1775e-06 | 9.2857e-08 | 2.291e-08 | 1.8705e-08 | 10 | 2 | 0.2 | 10 | 1.0669 | -8.674 | -2.2075 |  | 0.042948 | 0.035065 | 0.00087037 | 0.18746 | 0.18746 | Do a 1% point. 0.2 ohm in series with the Zetka.
23
24 |  |  | The error is an order of magnitude higher then previously, check the internal wireing in the 2 stage as it was disturbed when the outer primary was repaired.
25 |  |  | Moving the middle primary wires changes the error for X and Y. | [6 empty] | Found that the Aux Secondary live lead solder joint was cracked from solder tag internally. (high resistance joint).
26
27 |  |  | Shunt | X | Y | Sh V | X stdev | Y stdev | SR830 | SR830 | [6 empty] | SD ppm | SD ppm | SD ppm | noise sd | noise sd
28 |  |  | Volts | V | V | stdev/100 | V | V | Gain | Reserve | Shunt | Com R | % load | X ppm | Y ppm |  | X | Y | shunt V | X | Y
29 |  | 53 Hz zero burden on aux sec. (10 ohm common resistor). | 0.10424 | -3.9192e-06 | -1.1199e-06 | 1.5112e-08 | 1.3463e-08 | 1.1163e-08 | 10 | 2 | 0.2 | 10 | 10.4243 | -0.75193 | -0.21487 |  | 0.002583 | 0.0021418 | 1.4497e-05 | 0.019186 | 0.019186 | Initial Setup
30 |  |  | 0.10424 | 0.0013694 | 4.4048e-05 | 1.1255e-08 | 7.7175e-08 | 5.96e-08 | 18 | 2 | 0.2 | 10 | 10.4237 | 262.7419 | 8.4516 |  | 0.014808 | 0.011436 | 1.0798e-05 | 0.019187 | 0.019187 | 1k ohm across main secondary
31 |  |  | 0.10419 | -1.7607e-05 | 0.0004583 | 5.3524e-08 | 2.23e-08 | 2.414e-08 | 16 | 2 | 0.2 | 10 | 10.4195 | -3.3797 | 87.9704 |  | 0.0042805 | 0.0046337 | 5.1369e-05 | 0.019195 | 0.019195 | 1uF across main secondary
32 |  |  | 1.2493 | -1.9923e-05 | -1.6144e-05 | 4.4107e-06 | 2.1812e-08 | 4.8358e-08 | 13 | 2 | 0.2 | 10 | 124.931 | -0.31894 | -0.25844 |  | 0.00034918 | 0.00077415 | 0.00035305 | 0.0016009 | 0.0016009 | Start at 125%. 3458A on 12V range.
33 |  |  | 1.1996 | -1.9693e-05 | -1.5639e-05 | 2.3876e-06 | 1.0573e-07 | 6.1859e-08 | 13 | 2 | 0.2 | 10 | 119.9649 | -0.32832 | -0.26073 |  | 0.0017628 | 0.0010313 | 0.00019903 | 0.0016672 | 0.0016672 | 120%. 3458A at 1.2V range
34 |  |  | 1.0006 | -1.8196e-05 | -1.319e-05 | 4.5239e-07 | 5.919e-08 | 6.9388e-08 | 13 | 2 | 0.2 | 10 | 100.06 | -0.3637 | -0.26365 |  | 0.0011831 | 0.0013869 | 4.5212e-05 | 0.0019988 | 0.0019988
35 |  |  | 0.60026 | -1.4699e-05 | -7.3077e-06 | 6.3259e-07 | 3.3027e-08 | 3.0845e-08 | 12 | 2 | 0.2 | 10 | 60.0261 | -0.48976 | -0.24348 |  | 0.0011004 | 0.0010277 | 0.00010539 | 0.0033319 | 0.0033319
36 |  |  | 0.40036 | -1.1507e-05 | -4.6593e-06 | 4.1582e-07 | 4.1492e-08 | 4.4619e-08 | 12 | 2 | 0.2 | 10 | 40.0361 | -0.57481 | -0.23276 |  | 0.0020727 | 0.0022289 | 0.00010386 | 0.0049955 | 0.0049955
37 |  |  | 0.20093 | -6.891e-06 | -2.2029e-06 | 1.7276e-07 | 5.2078e-08 | 8.0782e-08 | 11 | 2 | 0.2 | 10 | 20.0926 | -0.68593 | -0.21927 |  | 0.0051838 | 0.008041 | 8.598e-05 | 0.0099539 | 0.0099539
38 |  |  | 0.1008 | -3.8076e-06 | -1.1252e-06 | 6.8879e-08 | 2.0816e-08 | 1.8181e-08 | 10 | 2 | 0.2 | 10 | 10.08 | -0.75549 | -0.22326 |  | 0.0041302 | 0.0036073 | 6.8333e-05 | 0.019841 | 0.019841
39 |  |  | 0.050013 | -2.0001e-06 | -5.537e-07 | 4.3853e-08 | 1.4684e-08 | 1.5406e-08 | 10 | 2 | 0.2 | 10 | 5.0013 | -0.79981 | -0.22142 |  | 0.0058722 | 0.0061608 | 8.7681e-05 | 0.039989 | 0.039989
40 |  |  | 0.010249 | -4.5818e-07 | -9.8404e-08 | 7.901e-08 | 6.8144e-08 | 6.2359e-08 | 8 | 2 | 0.2 | 10 | 1.0249 | -0.89413 | -0.19203 |  | 0.13298 | 0.12169 | 0.00077093 | 0.19515 | 0.19515 | Do a 1% point. 0.2 ohm in series with the Zetka.
41
42 |  | 47 Hz zero burden on aux sec. (10 ohm common resistor). | 1.2499 | -2.3253e-05 | -2.0335e-05 | 2.5367e-06 | 2.2271e-08 | 2.0776e-08 | 13 | 2 | 0.2 | 10 | 124.989 | -0.37208 | -0.32539 |  | 0.00035637 | 0.00033245 | 0.00020295 | 0.0016001 | 0.0016001 | Start at 125%. 3458A on 12V range.
43 |  |  | 1.1997 | -2.2936e-05 | -1.9628e-05 | 1.8788e-06 | 2.6818e-08 | 2.3449e-08 | 13 | 2 | 0.2 | 10 | 119.9693 | -0.38237 | -0.32721 |  | 0.00044709 | 0.00039091 | 0.0001566 | 0.0016671 | 0.0016671 | 120%. 3458A at 1.2V range
44 |  |  | 1.0009 | -2.1138e-05 | -1.6689e-05 | 6.8904e-08 | 3.2168e-08 | 2.9301e-08 | 13 | 2 | 0.2 | 10 | 100.0884 | -0.42239 | -0.33349 |  | 0.00064279 | 0.0005855 | 6.8843e-06 | 0.0019982 | 0.0019982
45 |  |  | 0.60035 | -1.7344e-05 | -9.6645e-06 | 6.9889e-07 | 5.1956e-08 | 3.3613e-08 | 12 | 2 | 0.2 | 10 | 60.0346 | -0.5778 | -0.32196 |  | 0.0017309 | 0.0011198 | 0.00011641 | 0.0033314 | 0.0033314
46 |  |  | 0.39995 | -1.392e-05 | -6.1039e-06 | 3.9673e-07 | 3.3555e-08 | 3.7474e-08 | 12 | 2 | 0.2 | 10 | 39.9948 | -0.69611 | -0.30524 |  | 0.001678 | 0.0018739 | 9.9195e-05 | 0.0050006 | 0.0050006
47 |  |  | 0.19988 | -8.4081e-06 | -2.9953e-06 | 1.9409e-07 | 4.9147e-08 | 4.6822e-08 | 11 | 2 | 0.2 | 10 | 19.9879 | -0.84131 | -0.29971 |  | 0.0049177 | 0.004685 | 9.7103e-05 | 0.010006 | 0.010006
48 |  |  | 0.10045 | -4.7447e-06 | -1.5007e-06 | 7.0499e-08 | 4.1624e-08 | 2.1782e-08 | 10 | 2 | 0.2 | 10 | 10.0454 | -0.94464 | -0.29878 |  | 0.0082871 | 0.0043367 | 7.0181e-05 | 0.01991 | 0.01991
49 |  |  | 0.050658 | -2.5761e-06 | -7.2535e-07 | 3.1498e-08 | 1.7899e-08 | 2.2677e-08 | 10 | 2 | 0.2 | 10 | 5.0658 | -1.0171 | -0.28637 |  | 0.0070665 | 0.0089531 | 6.2178e-05 | 0.03948 | 0.03948
50 |  |  | 0.010222 | -5.561e-07 | -1.4667e-07 | 6.584e-08 | 3.6811e-08 | 2.7396e-08 | 8 | 2 | 0.2 | 10 | 1.0222 | -1.0881 | -0.28698 |  | 0.072025 | 0.053605 | 0.00064412 | 0.19566 | 0.19566 | Do a 1% point. 0.2 ohm in series with the Zetka.
51
52 | 2018-08-01 |  | The error returned to mV. Found that the 1.7mm solid Cu wire to the Low (black) terminal was disconnected and the terminal was loose. Replaced the terminal and resoldered the wire to a new solder lug.
53 |  |  | Repeat the Self Cal
54 |  |  | Shunt | X | Y | Sh V | X stdev | Y stdev | SR830 | SR830 | [6 empty] | SD ppm | SD ppm | SD ppm | noise sd | noise sd
55 |  |  | Volts | V | V | stdev/100 | V | V | Gain | Reserve | Shunt | Com R | % load | X ppm | Y ppm |  | X | Y | shunt V | X | Y
56 |  | 53 Hz zero burden on aux sec. (10 ohm common resistor). | 0.10073 | -2.5618e-06 | -3.8222e-07 | 2.953e-08 | 2.2198e-08 | 2.6252e-08 | 10 | 2 | 0.2 | 10 | 10.0732 | -0.50863 | -0.075888 |  | 0.0044072 | 0.0052123 | 2.9316e-05 | 0.019855 | 0.019855 | Initial Setup
57 |  |  | 0.1006 | 0.0013465 | 4.1795e-05 | 1.4762e-08 | 4.8e-08 | 5.8396e-08 | 18 | 2 | 0.2 | 10 | 10.06 | 267.6839 | 8.3091 |  | 0.0095427 | 0.011609 | 1.4674e-05 | 0.019881 | 0.019881 | 1k ohm across main secondary
58 |  |  | 0.10064 | -1.5289e-05 | 0.00044999 | 3.0354e-08 | 5.1615e-08 | 4.76e-08 | 18 | 2 | 0.2 | 10 | 10.0635 | -3.0384 | 89.4303 |  | 0.010258 | 0.0094599 | 3.0163e-05 | 0.019874 | 0.019874 | 1uF across main secondary
59 |  |  | 1.2493 | -1.3317e-05 | -6.2982e-06 | 4.8677e-06 | 5.1997e-08 | 2.0837e-08 | 12 | 2 | 0.2 | 10 | 124.9328 | -0.21318 | -0.10083 |  | 0.0008324 | 0.00033357 | 0.00038962 | 0.0016009 | 0.0016009 | Start at 125%. 3458A on 12V range.
60 |  |  | 1.1992 | -1.3003e-05 | -6.1099e-06 | 2.5599e-06 | 2.9645e-08 | 2.703e-08 | 12 | 2 | 0.2 | 10 | 119.9205 | -0.21686 | -0.1019 |  | 0.00049442 | 0.0004508 | 0.00021347 | 0.0016678 | 0.0016678 | 120%. 3458A at 1.2V range
61 |  |  | 1.0026 | -1.1927e-05 | -5.215e-06 | 4.0252e-07 | 2.5446e-08 | 2.0453e-08 | 12 | 2 | 0.2 | 10 | 100.2644 | -0.2379 | -0.10402 |  | 0.00050757 | 0.00040799 | 4.0145e-05 | 0.0019947 | 0.0019947
62 |  |  | 0.60131 | -9.6089e-06 | -2.7846e-06 | 5.4873e-07 | 2.8477e-08 | 2.6778e-08 | 11 | 2 | 0.2 | 10 | 60.1312 | -0.3196 | -0.092616 |  | 0.00094718 | 0.00089066 | 9.1255e-05 | 0.0033261 | 0.0033261
63 |  |  | 0.40028 | -7.5746e-06 | -1.6338e-06 | 3.8545e-07 | 1.5527e-08 | 1.4833e-08 | 11 | 2 | 0.2 | 10 | 40.0279 | -0.37847 | -0.081632 |  | 0.00077581 | 0.00074113 | 9.6295e-05 | 0.0049965 | 0.0049965
64 |  |  | 0.19991 | -4.4998e-06 | -7.5213e-07 | 1.7214e-07 | 3.1157e-08 | 2.4658e-08 | 10 | 2 | 0.2 | 10 | 19.9906 | -0.45019 | -0.075248 |  | 0.0031172 | 0.0024669 | 8.6112e-05 | 0.010005 | 0.010005
65 |  |  | 0.10052 | -2.4974e-06 | -3.4019e-07 | 7.3131e-08 | 2.1576e-08 | 2.1323e-08 | 10 | 2 | 0.2 | 10 | 10.0523 | -0.49688 | -0.067684 |  | 0.0042927 | 0.0042424 | 7.275e-05 | 0.019896 | 0.019896
66 |  |  | 0.050294 | -1.3771e-06 | -1.577e-07 | 4.7361e-08 | 2.0762e-08 | 1.7683e-08 | 9 | 2 | 0.2 | 10 | 5.0294 | -0.54764 | -0.062711 |  | 0.0082565 | 0.007032 | 9.4169e-05 | 0.039767 | 0.039767
67 |  |  | 0.010335 | -3.0099e-07 | -3.7803e-08 | 5.2626e-08 | 2.222e-08 | 1.7361e-08 | 7 | 2 | 0.2 | 10 | 1.0335 | -0.58247 | -0.073157 |  | 0.043 | 0.033597 | 0.00050921 | 0.19352 | 0.19352 | Do a 1% point. 0.2 ohm in series with the Zetka.
68
69 |  | 47 Hz zero burden on aux sec. (10 ohm common resistor). | 1.2514 | -1.5265e-05 | -9.1327e-06 | 3.5766e-06 | 3.3886e-08 | 2.6752e-08 | 12 | 2 | 0.2 | 10 | 125.1351 | -0.24397 | -0.14596 |  | 0.00054159 | 0.00042756 | 0.00028582 | 0.0015983 | 0.0015983 | Start at 125%. 3458A on 12V range.
70 |  |  | 1.1989 | -1.4778e-05 | -8.8e-06 | 1.7555e-06 | 2.0144e-08 | 2.7037e-08 | 12 | 2 | 0.2 | 10 | 119.8908 | -0.24653 | -0.1468 |  | 0.00033604 | 0.00045103 | 0.00014643 | 0.0016682 | 0.0016682 | 120%. 3458A at 1.2V range
71 |  |  | 1.0001 | -1.3542e-05 | -7.5639e-06 | 1.0321e-07 | 1.4892e-08 | 1.3209e-08 | 12 | 2 | 0.2 | 10 | 100.0052 | -0.27082 | -0.15127 |  | 0.00029782 | 0.00026416 | 1.032e-05 | 0.0019999 | 0.0019999
72 |  |  | 0.59973 | -1.1156e-05 | -4.3445e-06 | 5.6221e-07 | 2.4127e-08 | 4.3327e-08 | 12 | 2 | 0.2 | 10 | 59.9727 | -0.37204 | -0.14488 |  | 0.0008046 | 0.0014449 | 9.3744e-05 | 0.0033348 | 0.0033348
73 |  |  | 0.40064 | -8.9943e-06 | -2.6978e-06 | 3.0093e-07 | 2.7015e-08 | 4.0438e-08 | 11 | 2 | 0.2 | 10 | 40.0636 | -0.449 | -0.13467 |  | 0.0013486 | 0.0020187 | 7.5114e-05 | 0.0049921 | 0.0049921
74 |  |  | 0.20025 | -5.4486e-06 | -1.3189e-06 | 1.6609e-07 | 3.3877e-08 | 3.528e-08 | 11 | 2 | 0.2 | 10 | 20.0246 | -0.54419 | -0.13173 |  | 0.0033835 | 0.0035236 | 8.2942e-05 | 0.0099877 | 0.0099877
75 |  |  | 0.10055 | -3.0696e-06 | -6.5086e-07 | 9.4645e-08 | 2.848e-08 | 1.5685e-08 | 10 | 2 | 0.2 | 10 | 10.0546 | -0.61059 | -0.12947 |  | 0.0056651 | 0.0031199 | 9.4131e-05 | 0.019891 | 0.019891
76 |  |  | 0.049978 | -1.6346e-06 | -3.2598e-07 | 1.5089e-08 | 1.7264e-08 | 2.7856e-08 | 9 | 2 | 0.2 | 10 | 4.9978 | -0.65413 | -0.13045 |  | 0.0069085 | 0.011147 | 3.0191e-05 | 0.040018 | 0.040018
77 |  |  | 0.010228 | -3.5723e-07 | -6.9596e-08 | 4.3573e-08 | 4.2104e-08 | 3.515e-08 | 7 | 2 | 0.2 | 10 | 1.0228 | -0.69852 | -0.13609 |  | 0.082329 | 0.068732 | 0.00042601 | 0.19554 | 0.19554 | Do a 1% point. 0.2 ohm in series with the Zetka.
78
79 |  | Calibration of 20:5 ratio of Tb vs inner pri 20:5 ratio of Ta | [5 empty] | unmarked terminal of Ta pri wagnered to earth with IVD, phase SR830 to MSL shunt.
80 |  |  | Shunt | X | Y | Sh V | X stdev | Y stdev | [8 empty] | SD ppm | SD ppm | SD ppm | noise sd | noise sd
81 |  |  | Volts | V | V | stdev/100 | V | V |  |  | Shunt | Com R | % load | X ppm | Y ppm |  | X | Y | shunt V | X | Y
82 |  | 53Hz | 0.61864 | 8.2021e-07 | -7.4906e-07 | 1.8094e-06 | 2.3166e-07 | 3.3631e-08 | 9 | 2 | 0.2 | 10 | 61.8644 | 0.026517 | -0.024216 |  | 0.0074894 | 0.0010872 | 0.00029248 | 0.0032329 | 0.0032329 | Initial Setup
83 |  |  | 0.61804 | 0.0080883 | 0.00044074 | 2.8726e-06 | 7.0246e-07 | 3.0543e-07 | 20 | 2 | 0.2 | 10 | 61.8042 | 261.7383 | 14.2625 |  | 0.022732 | 0.0098838 | 0.00046479 | 0.003236 | 0.003236 | 1k ohm across Ta main secondary
84 |  |  | 0.61723 | -0.00013805 | 0.0026958 | 1.1089e-06 | 9.9727e-07 | 1.0282e-06 | 19 | 2 | 0.2 | 10 | 61.7225 | -4.4731 | 87.3526 |  | 0.032315 | 0.033318 | 0.00017965 | 0.0032403 | 0.0032403 | 1uF across Ta main secondary
85 |  |  | 1.2482 | -3.3808e-06 | -2.4216e-06 | 2.2481e-05 | 5.3475e-08 | 1.2895e-08 | 10 | 2 | 0.2 | 10 | 124.8242 | -0.054168 | -0.038801 |  | 0.00085681 | 0.00020661 | 0.001801 | 0.0016023 | 0.0016023 | Start at 125%. 3458A on 12V range.
86 |  |  | 1.1952 | -4.5898e-06 | -2.1937e-06 | 1.1594e-05 | 1.2595e-08 | 1.9463e-08 | 10 | 2 | 0.2 | 10 | 119.5181 | -0.076805 | -0.036709 |  | 0.00021076 | 0.0003257 | 0.00097004 | 0.0016734 | 0.0016734 | 120%. 3458A at 1.2V range
87 |  |  | 1.0031 | -3.0076e-06 | -1.9507e-06 | 9.9994e-07 | 1.116e-07 | 2.2015e-08 | 10 | 2 | 0.2 | 10 | 100.3091 | -0.059966 | -0.038893 |  | 0.0022252 | 0.00043894 | 9.9686e-05 | 0.0019938 | 0.0019938
88 |  |  | 0.60285 | -1.9597e-06 | -9.312e-07 | 6.121e-06 | 1.6915e-08 | 1.2981e-08 | 9 | 2 | 0.2 | 10 | 60.2852 | -0.065014 | -0.030893 |  | 0.00056115 | 0.00043065 | 0.0010153 | 0.0033176 | 0.0033176
89 |  |  | 0.40084 | -1.2958e-06 | -8.0334e-07 | 3.639e-06 | 2.0607e-08 | 1.7997e-08 | 9 | 2 | 0.2 | 10 | 40.0841 | -0.064653 | -0.040083 |  | 0.0010282 | 0.00089795 | 0.00090783 | 0.0049895 | 0.0049895
90 |  |  | 0.20039 | -7.3284e-07 | -4.6212e-07 | 1.2216e-06 | 1.52e-08 | 2.4215e-08 | 9 | 2 | 0.2 | 10 | 20.0391 | -0.073141 | -0.046122 |  | 0.001517 | 0.0024167 | 0.0006096 | 0.0099805 | 0.0099805
91 |  |  | 0.1001 | -4.2458e-07 | -2.1018e-07 | 3.9408e-07 | 5.601e-08 | 5.363e-08 | 8 | 2 | 0.2 | 10 | 10.01 | -0.084832 | -0.041995 |  | 0.011191 | 0.010715 | 0.00039368 | 0.01998 | 0.01998
92 |  |  | 0.049845 | -2.1096e-07 | -1.1895e-07 | 6.1912e-07 | 2.4698e-08 | 1.3037e-08 | 8 | 2 | 0.2 | 10 | 4.9845 | -0.084647 | -0.047726 |  | 0.00991 | 0.005231 | 0.0012421 | 0.040124 | 0.040124
93 |  |  | 0.0099378 | -5.1147e-08 | -3.3778e-08 | 1.0283e-07 | 1.0896e-08 | 1.7821e-08 | 8 | 2 | 0.2 | 10 | 0.99378 | -0.10293 | -0.067979 |  | 0.021929 | 0.035865 | 0.0010348 | 0.20125 | 0.20125 | Do a 1% point. 1 ohm in series with the Zetka.
94
95 |  |  | 0.60036 | -0.010094 | -0.00058961 | 2.3727e-06 | 1.1269e-05 | 1.0821e-06 | 20 | 2 | 0.2 | 10 | 60.036 | -336.2655 | -19.642 |  | 0.37541 | 0.036047 | 0.0003952 | 0.0033313 | 0.0033313 | 1k ohm across Tb main secondary
96 |  |  | 0.59778 | 0.00018588 | -0.0033369 | 5.774e-07 | 1.1352e-07 | 6.2601e-07 | 19 | 2 | 0.2 | 10 | 59.7777 | 6.2192 | -111.6423 |  | 0.003798 | 0.020944 | 9.6592e-05 | 0.0033457 | 0.0033457 | 1uF across Tb main secondary
97
98
99 |  | Calibration of 100:5 ratio, middle primary of Ta, vs 100:5 ratio, inner primary of Tb | [8 empty] | Wagner at Tb end of link, phase SR830 to the MSL shunt.
100 |  |  | 0.10086 | 1.1862e-06 | 5.4529e-07 | 1.5166e-08 | 2.3049e-08 | 2.4331e-08 | 9 | 2 | 0.2 | 10 | 10.0861 | 0.23522 | 0.10813 |  | 0.0045705 | 0.0048246 | 1.5037e-05 |  |  | Initial Setup
101 |  |  | 0.10072 | -0.00019172 | 1.4134e-06 | 5.2874e-08 | 3.2062e-08 | 1.9865e-08 | 15 | 2 | 0.2 | 10 | 10.0721 | -38.0701 | 0.28065 |  | 0.0063664 | 0.0039446 | 5.2495e-05 |  |  | 10k ohm across Ta main secondary
102 |  |  | 0.10105 | -3.8356e-06 | -0.00059638 | 7.5039e-08 | 2.8427e-08 | 1.7831e-07 | 17 | 2 | 0.2 | 10 | 10.105 | -0.75915 | -118.0367 |  | 0.0056264 | 0.035292 | 7.4259e-05 |  |  | 1uF across Ta main secondary
103 |  |  | 1.2485 | 5.9921e-06 | 5.7682e-06 | 1.0546e-05 | 2.4287e-07 | 2.8207e-08 | 11 | 2 | 0.2 | 10 | 124.8517 | 0.095988 | 0.092401 |  | 0.0038905 | 0.00045184 | 0.00084466 |  | Zetka clipping (Red) | Start at 125%. 3458A on 12V range.
104 |  |  | 1.1985 | 1.0234e-05 | 6.0902e-06 | 6.3268e-06 | 4.8872e-07 | 1.9692e-08 | 12 | 2 | 0.2 | 10 | 119.8509 | 0.17078 | 0.10163 |  | 0.0081555 | 0.0003286 | 0.00052789 |  |  | 120%. 3458A at 1.2V range
105 |  |  | 1.0007 | 8.3128e-06 | 5.4102e-06 | 1.3874e-06 | 1.616e-08 | 2.1501e-08 | 11 | 2 | 0.2 | 10 | 100.0691 | 0.16614 | 0.10813 |  | 0.00032298 | 0.00042972 | 0.00013865
106 |  |  | 0.60042 | 5.1426e-06 | 3.6261e-06 | 1.2798e-06 | 3.0151e-08 | 3.1332e-08 | 11 | 2 | 0.2 | 10 | 60.0415 | 0.1713 | 0.12079 |  | 0.0010043 | 0.0010437 | 0.00021316
107 |  |  | 0.40065 | 4.0755e-06 | 2.2477e-06 | 1.5912e-06 | 6.6251e-08 | 5.5817e-08 | 10 | 2 | 0.2 | 10 | 40.0646 | 0.20345 | 0.11221 |  | 0.0033072 | 0.0027863 | 0.00039716
108 |  |  | 0.20045 | 2.3414e-06 | 1.08e-06 | 5.9081e-07 | 3.3123e-08 | 2.2056e-08 | 10 | 2 | 0.2 | 10 | 20.0454 | 0.23361 | 0.10776 |  | 0.0033047 | 0.0022006 | 0.00029473
109 |  |  | 0.10024 | 1.2858e-06 | 5.399e-07 | 7.5468e-08 | 1.7755e-08 | 2.3779e-08 | 9 | 2 | 0.2 | 10 | 10.0238 | 0.25655 | 0.10772 |  | 0.0035426 | 0.0047445 | 7.5289e-05
110 |  |  | 0.05004 | 6.8873e-07 | 2.6071e-07 | 7.7413e-08 | 3.6679e-08 | 1.8794e-08 | 8 | 2 | 0.2 | 10 | 5.004 | 0.27527 | 0.1042 |  | 0.01466 | 0.0075116 | 0.0001547
111 |  |  | 0.010096 | 1.3753e-07 | 7.3211e-08 | 7.5344e-08 | 1.9976e-08 | 1.3989e-08 | 7 | 2 | 0.2 | 10 | 1.0096 | 0.27244 | 0.14503 |  | 0.039574 | 0.027712 | 0.0007463 |  |  | Do a 1% point. No additional resistance needed.
112
113
114 | 2018-08-02 | Calibration of 100:5 ratio, middle primary of Ta, vs 100:5 ratio, inner primary of Tb | [8 empty] | Wagner at Tb end of link, phase SR830 to the MSL shunt.
115 |  | Repeat last point | 1.2536 | 8.649e-06 | 6.986e-06 | 7.1655e-06 | 1.8443e-07 | 2.7369e-08 | 11 | 2 | 0.2 | 10 | 125.3561 | 0.13799 | 0.11146 |  | 0.0029425 | 0.00043666 | 0.00057161 |  | Zetka clipping (Red)
116 |  |  | 1.1979 | 6.3911e-06 | 6.6023e-06 | 7.7014e-06 | 2.1237e-07 | 4.2278e-08 | 11 | 2 | 0.2 | 10 | 119.7872 | 0.10671 | 0.11023 |  | 0.0035458 | 0.00070589 | 0.00064292
117 |  |  | 1.001 | 5.4642e-06 | 5.7776e-06 | 5.9651e-06 | 7.5163e-08 | 4.2117e-08 | 11 | 2 | 0.2 | 10 | 100.0955 | 0.10918 | 0.11544 |  | 0.0015018 | 0.00084153 | 0.00059594
118 |  |  | 0.60472 | 4.5956e-06 | 3.8326e-06 | 3.7824e-07 | 8.2363e-08 | 6.3389e-08 | 11 | 2 | 0.2 | 10 | 60.4718 | 0.15199 | 0.12676 |  | 0.002724 | 0.0020965 | 6.2548e-05
119 |  |  | 0.40023 | 3.7118e-06 | 2.4716e-06 | 1.1904e-06 | 1.4078e-08 | 1.8475e-08 | 10 | 2 | 0.2 | 10 | 40.0229 | 0.18548 | 0.12351 |  | 0.00070349 | 0.00092321 | 0.00029744
120 |  |  | 0.20032 | 2.2019e-06 | 1.251e-06 | 5.3954e-07 | 1.1222e-08 | 1.7871e-08 | 10 | 2 | 0.2 | 10 | 20.0318 | 0.21984 | 0.1249 |  | 0.0011205 | 0.0017843 | 0.00026934
121 |  |  | 0.10027 | 1.2314e-06 | 6.1348e-07 | 2.0371e-07 | 2.634e-08 | 3.1253e-08 | 9 | 2 | 0.2 | 10 | 10.0271 | 0.24561 | 0.12236 |  | 0.0052537 | 0.0062337 | 0.00020316
122 |  |  | 0.050571 | 6.6001e-07 | 3.2128e-07 | 8.1722e-08 | 1.7995e-08 | 2.2417e-08 | 8 | 2 | 0.2 | 10 | 5.0571 | 0.26103 | 0.12706 |  | 0.0071169 | 0.0088656 | 0.0001616
123 |  |  | 0.010051 | 1.4303e-07 | 5.904e-08 | 3.3117e-08 | 1.9983e-08 | 1.6336e-08 | 7 | 2 | 0.2 | 10 | 1.0051 | 0.28463 | 0.11749 |  | 0.039765 | 0.032508 | 0.0003295 |  |  | Do a 1% point. No additional resistance needed.
124
125 |  | Calibration of 100:5 ratio, outer primary of Tb, vs 100:5 ratio, middle primary of Ta | [8 empty] | Wagner at Tb end of link, phase SR830 to the MSL shunt.
126 |  |  | 0.10478 | -2.1552e-07 | -8.7692e-08 | 1.7123e-08 | 1.6911e-08 | 1.1222e-08 | 7 | 2 | 0.2 | 10 | 10.4783 | -0.041136 | -0.016738 |  | 0.0032277 | 0.0021419 | 1.6342e-05 |  |  | Initial Setup
127 |  |  | 0.10442 | -0.00018463 | -7.5843e-06 | 5.5102e-08 | 1.614e-08 | 1.9352e-08 | 15 | 2 | 0.2 | 10 | 10.4418 | -35.3635 | -1.4527 |  | 0.0030913 | 0.0037066 | 5.2771e-05 |  |  | 10k ohm across Tb main secondary
128 |  |  | 0.10563 | 2.1419e-05 | -0.00060875 | 2.1433e-07 | 3.8969e-08 | 1.038e-07 | 17 | 2 | 0.2 | 10 | 10.5628 | 4.0555 | -115.2635 |  | 0.0073784 | 0.019654 | 0.00020291 |  |  | 1uF across Tb main secondary
129 |  |  | 1.2513 | -9.9664e-07 | -1.0917e-06 | 1.0376e-05 | 6.3058e-08 | 1.8242e-08 | 9 | 2 | 0.2 | 10 | 125.1251 | -0.01593 | -0.017449 |  | 0.0010079 | 0.00029158 | 0.00082922 |  |  | Start at 125%. 3458A on 12V range.
130 |  |  | 1.198 | -1.1732e-06 | -1.0612e-06 | 5.7854e-06 | 1.803e-08 | 1.925e-08 | 9 | 2 | 0.2 | 10 | 119.798 | -0.019587 | -0.017716 |  | 0.00030101 | 0.00032138 | 0.00048293 |  |  | 120%. 3458A at 1.2V range
131 |  |  | 1.0011 | -1.5706e-06 | -8.8567e-07 | 6.3748e-07 | 1.7877e-08 | 1.3466e-08 | 9 | 2 | 0.2 | 10 | 100.1098 | -0.031378 | -0.017694 |  | 0.00035714 | 0.00026903 | 6.3678e-05
132 |  |  | 0.60017 | -1.4923e-06 | -4.2484e-07 | 1.0631e-06 | 1.0823e-08 | 2.6198e-08 | 9 | 2 | 0.2 | 10 | 60.0168 | -0.049731 | -0.014157 |  | 0.00036066 | 0.00087303 | 0.00017713
133 |  |  | 0.40012 | -1.1169e-06 | -3.6244e-07 | 7.2048e-07 | 9.8564e-09 | 1.8892e-08 | 9 | 2 | 0.2 | 10 | 40.0116 | -0.055829 | -0.018117 |  | 0.00049268 | 0.00094432 | 0.00018007
134 |  |  | 0.19997 | -6.7534e-07 | -2.2279e-07 | 4.561e-07 | 1.5968e-08 | 1.7202e-08 | 8 | 2 | 0.2 | 10 | 19.9968 | -0.067545 | -0.022282 |  | 0.001597 | 0.0017204 | 0.00022809
135 |  |  | 0.10037 | -3.7501e-07 | -1.3897e-07 | 1.9686e-07 | 1.39e-08 | 1.4283e-08 | 8 | 2 | 0.2 | 10 | 10.0371 | -0.074725 | -0.027692 |  | 0.0027698 | 0.002846 | 0.00019613
136 |  |  | 0.10052 | -3.6226e-07 | -1.2849e-07 | 1.451e-07 | 2.4981e-08 | 1.0552e-08 | 8 | 2 | 0.2 | 10 | 10.0518 | -0.072079 | -0.025566 |  | 0.0049705 | 0.0020995 | 0.00014435 |  |  | Additional 10% points
137 |  |  | 0.099847 | -3.706e-07 | -9.0883e-08 | 1.6334e-07 | 1.9176e-08 | 1.9226e-08 | 8 | 2 | 0.2 | 10 | 9.9847 | -0.074234 | -0.018204 |  | 0.003841 | 0.0038511 | 0.00016359 |  |  | Additional 10% points
138 |  |  | 0.049833 | -2.0304e-07 | -8.6512e-08 | 6.5881e-08 | 2.1184e-08 | 2.1883e-08 | 8 | 2 | 0.2 | 10 | 4.9833 | -0.081487 | -0.034721 |  | 0.0085021 | 0.0087826 | 0.0001322
139 |  |  | 0.049867 | -1.9732e-07 | -5.1665e-08 | 6.7288e-08 | 2.5941e-08 | 2.2409e-08 | 8 | 2 | 0.2 | 10 | 4.9867 | -0.079139 | -0.020721 |  | 0.010404 | 0.0089873 | 0.00013493 |  |  | Additional 5% points
140 |  |  | 0.0499 | -1.9916e-07 | -5.4081e-08 | 6.4394e-08 | 1.7959e-08 | 1.7726e-08 | 8 | 2 | 0.2 | 10 | 4.99 | -0.079826 | -0.021676 |  | 0.0071982 | 0.0071046 | 0.00012905 |  |  | Additional 5% points
141 |  |  | 0.049958 | -1.7751e-07 | -7.9006e-08 | 6.7035e-08 | 7.5659e-08 | 5.5737e-08 | 8 | 2 | 0.2 | 10 | 4.9958 | -0.071064 | -0.031629 |  | 0.030289 | 0.022314 | 0.00013418 |  |  | Additional 5% points
142 |  |  | 0.050006 | -1.9411e-07 | -8.5256e-08 | 5.3652e-08 | 7.0916e-08 | 6.5999e-08 | 7 | 2 | 0.2 | 10 | 5.0006 | -0.077635 | -0.034098 |  | 0.028363 | 0.026396 | 0.00010729 |  |  | Additional 5% points
143 |  |  | 0.010013 | -3.8433e-08 | -2.8534e-08 | 7.3394e-08 | 2.1625e-08 | 3.7432e-08 | 7 | 2 | 0.2 | 10 | 1.0013 | -0.07677 | -0.056995 |  | 0.043195 | 0.07477 | 0.00073302 |  |  | Do a 1% point. No additional resistance needed.
144 |  |  | 0.010034 | -2.6908e-08 | 2.8958e-09 | 2.8792e-08 | 2.7437e-08 | 2.5167e-08 | 7 | 2 | 0.2 | 10 | 1.0034 | -0.053636 | 0.0057723 |  | 0.05469 | 0.050165 | 0.00028696 |  |  | Additional 1% points
145 |  |  | 0.010022 | -5.7745e-08 | -1.7865e-08 | 1.012e-07 | 2.9512e-08 | 2.9139e-08 | 7 | 2 | 0.2 | 10 | 1.0022 | -0.11524 | -0.035652 |  | 0.058894 | 0.05815 | 0.0010097 |  |  | Additional 1% points
146 |  |  | 0.01004 | -5.3751e-08 | -7.0548e-09 | 8.8608e-08 | 4.7204e-08 | 3.3258e-08 | 7 | 2 | 0.2 | 10 | 1.004 | -0.10707 | -0.014053 |  | 0.094032 | 0.06625 | 0.00088254 |  |  | Additional 1% points
147 |  |  | 0.010031 | -4.031e-08 | 5.6316e-09 | 5.7969e-08 | 2.406e-08 | 2.0802e-08 | 7 | 2 | 0.2 | 10 | 1.0031 | -0.080371 | 0.011228 |  | 0.047972 | 0.041475 | 0.0005779 |  |  | Additional 1% points
148 |  |  | 0.010045 | -5.5979e-08 | -8.5333e-09 | 4.3181e-08 | 4.7496e-08 | 5.645e-08 | 7 | 2 | 0.2 | 10 | 1.0045 | -0.11146 | -0.01699 |  | 0.094567 | 0.1124 | 0.00042988 |  |  | Additional 1% points
149 |  |  | 0.010037 | -3.7645e-08 | -5.7291e-09 | 4.879e-08 | 2.8198e-08 | 1.9176e-08 | 7 | 2 | 0.2 | 10 | 1.0037 | -0.075013 | -0.011416 |  | 0.056189 | 0.038211 | 0.00048612 |  |  | Additional 1% points
150 |  |  | 0.010041 | -4.4749e-08 | -9.0106e-09 | 2.0761e-08 | 1.3271e-08 | 2.2271e-08 | 7 | 2 | 0.2 | 10 | 1.0041 | -0.089133 | -0.017948 |  | 0.026435 | 0.044361 | 0.00020677 |  |  | Additional 1% points
151 |  |  | 0.010064 | -3.6278e-08 | -5.9329e-09 | 1.0046e-07 | 1.8937e-08 | 2.564e-08 | 7 | 2 | 0.2 | 10 | 1.0064 | -0.072096 | -0.01179 |  | 0.037634 | 0.050955 | 0.00099826 |  |  | Additional 1% points
152 |  |  | 0.0099026 | -3.9023e-08 | -2.5847e-08 | 6.7076e-08 | 2.1796e-08 | 2.2097e-08 | 7 | 2 | 0.2 | 10 | 0.99026 | -0.078813 | -0.052203 |  | 0.044021 | 0.044629 | 0.00067736 |  |  | Additional 1% points
153 |  |  | 0.0098829 | -3.0703e-08 | -1.5213e-08 | 2.8596e-08 | 1.2422e-08 | 2.815e-08 | 7 | 2 | 0.2 | 10 | 0.98829 | -0.062134 | -0.030786 |  | 0.025139 | 0.056967 | 0.00028935 |  |  | Additional 1% points
154
155 |  | Calibration of 100:5 ratio, outer primary of Ta, vs 100:5 ratio, outer primary of Tb | [8 empty] | Wagner at Tb end of link, phase SR830 to the MSL shunt.
156 |  |  | 0.10037 | 1.3278e-06 | 3.7012e-07 | 6.7802e-09 | 6.3289e-08 | 8.0306e-08 | 9 | 2 | 0.2 | 10 | 10.0369 | 0.26457 | 0.073751 |  | 0.012611 | 0.016002 | 6.7553e-06 |  |  | Initial Setup
157 |  |  | 0.1002 | -0.00017815 | -8.9862e-06 | 3.4959e-08 | 6.9894e-08 | 6.482e-08 | 15 | 2 | 0.2 | 10 | 10.0201 | -35.5593 | -1.7936 |  | 0.013951 | 0.012938 | 3.4889e-05 |  |  | 10k ohm across Ta main secondary
158 |  |  | 0.099998 | 3.1349e-05 | -0.00061344 | 7.4213e-08 | 3.5134e-08 | 1.9e-07 | 17 | 2 | 0.2 | 10 | 9.9998 | 6.27 | -122.6912 |  | 0.0070269 | 0.038001 | 7.4215e-05 |  |  | 1uF across Ta main secondary
159 |  |  | 0.099769 | 1.4249e-06 | 4.8367e-07 | 9.2193e-08 | 8.6555e-08 | 7.6235e-08 | 9 | 2 | 0.2 | 10 | 9.9769 | 0.28564 | 0.096958 |  | 0.017351 | 0.015282 | 9.2406e-05
160 |  |  | 1.2499 | 1.0617e-05 | 4.8248e-06 | 8.2784e-06 | 8.8289e-08 | 1.9778e-08 | 12 | 2 | 0.2 | 10 | 124.9885 | 0.16989 | 0.077204 |  | 0.0014128 | 0.00031647 | 0.00066233 |  | Zetka clipping (Red) | Start at 125%. 3458A on 12V range.
161 |  |  | 1.244 | 1.0946e-05 | 4.8424e-06 | 8.4936e-06 | 1.8773e-07 | 1.1494e-07 | 12 | 2 | 0.2 | 10 | 124.4014 | 0.17599 | 0.077851 |  | 0.0030181 | 0.0018479 | 0.00068275 |  | Zetka clipping (Red) | 125%. 3458A on 12V range.
162 |  |  | 1.2486 | 1.8662e-05 | 5.9264e-06 | 1.1288e-05 | 2.3455e-06 | 5.4105e-07 | 12 | 2 | 0.2 | 10 | 124.8588 | 0.29894 | 0.09493 |  | 0.03757 | 0.0086666 | 0.00090407 |  | Zetka clipping (Red) | 125%. 3458A on 12V range.
163 |  |  | 1.2475 | 1.9172e-05 | 6.1008e-06 | 6.6031e-06 | 3.7974e-06 | 4.7212e-07 | 13 | 2 | 0.2 | 10 | 124.7492 | 0.30737 | 0.09781 |  | 0.06088 | 0.0075691 | 0.00052931 |  | Zetka clipping (Red) | 125%. 3458A on 12V range.
164 |  |  | 1.1988 | 1.1119e-05 | 5.1091e-06 | 4.8775e-06 | 3.3207e-07 | 2.165e-08 | 12 | 2 | 0.2 | 10 | 119.8778 | 0.1855 | 0.085238 |  | 0.0055402 | 0.0003612 | 0.00040687 |  | Zetka clipping (Red) | 125%. 3458A on 12V range.
165 |  |  | 1.1983 | 1.1127e-05 | 5.0826e-06 | 7.5549e-06 | 5.6332e-08 | 1.4931e-08 | 12 | 2 | 0.2 | 10 | 119.8308 | 0.18571 | 0.08483 |  | 0.00094019 | 0.00024921 | 0.00063046 |  | Zetka clipping (Red) | 120%. 3458A at 1.2V range
166 |  |  | 1.1986 | 1.0981e-05 | 5.1312e-06 | 5.8264e-06 | 3.4197e-08 | 2.3801e-08 | 12 | 2 | 0.2 | 10 | 119.8624 | 0.18323 | 0.085617 |  | 0.0005706 | 0.00039714 | 0.00048609 |  | Zetka clipping (Red)
167 |  |  | 1.1989 | 1.0774e-05 | 4.9922e-06 | 4.098e-06 | 2.8946e-08 | 1.7859e-08 | 12 | 2 | 0.2 | 10 | 119.894 | 0.17973 | 0.083277 |  | 0.00048286 | 0.00029791 | 0.0003418 |  | Zetka clipping (Red)
168 |  |  | 1.0008 | 9.6497e-06 | 4.3417e-06 | 4.3895e-07 | 3.0068e-08 | 1.7186e-08 | 11 | 2 | 0.2 | 10 | 100.0794 | 0.19284 | 0.086765 |  | 0.00060088 | 0.00034344 | 4.386e-05 |  | Zetka clipping (Red)
169 |  |  | 1.0005 | 9.496e-06 | 4.3195e-06 | 4.9285e-07 | 1.5818e-08 | 1.7026e-08 | 11 | 2 | 0.2 | 10 | 100.0488 | 0.18983 | 0.086348 |  | 0.0003162 | 0.00034036 | 4.9261e-05 |  | Zetka clipping (Red)
170 |  |  | 0.59958 | 6.534e-06 | 3.0132e-06 | 1.3428e-06 | 3.3314e-08 | 2.1347e-08 | 11 | 2 | 0.2 | 10 | 59.9578 | 0.21795 | 0.10051 |  | 0.0011113 | 0.00071206 | 0.00022396
171 |  |  | 0.60044 | 6.515e-06 | 2.9991e-06 | 1.3442e-06 | 2.0962e-08 | 2.3312e-08 | 11 | 2 | 0.2 | 10 | 60.0442 | 0.21701 | 0.099896 |  | 0.0006982 | 0.00077649 | 0.00022387
172 |  |  | 0.39964 | 5.0653e-06 | 1.8919e-06 | 1.1568e-06 | 3.0164e-08 | 2.669e-08 | 10 | 2 | 0.2 | 10 | 39.964 | 0.25349 | 0.094681 |  | 0.0015096 | 0.0013357 | 0.00028945
173 |  |  | 0.40021 | 5.0728e-06 | 1.9438e-06 | 1.0917e-06 | 7.8735e-08 | 6.96e-08 | 10 | 2 | 0.2 | 10 | 40.0212 | 0.25351 | 0.097136 |  | 0.0039347 | 0.0034781 | 0.00027278
174 |  |  | 0.1999 | 2.8775e-06 | 9.294e-07 | 5.5785e-07 | 7.1279e-08 | 5.4318e-08 | 10 | 2 | 0.2 | 10 | 19.9899 | 0.2879 | 0.092986 |  | 0.0071315 | 0.0054345 | 0.00027907
175 |  |  | 0.20021 | 2.9304e-06 | 9.3732e-07 | 4.9012e-07 | 1.9988e-08 | 2.2136e-08 | 10 | 2 | 0.2 | 10 | 20.0205 | 0.29274 | 0.093636 |  | 0.0019967 | 0.0022114 | 0.00024481
176 |  |  | 0.099972 | 1.6016e-06 | 4.0752e-07 | 2.4813e-07 | 2.4354e-08 | 2.1913e-08 | 9 | 2 | 0.2 | 10 | 9.9972 | 0.32042 | 0.081526 |  | 0.0048721 | 0.0043839 | 0.0002482
177 |  |  | 0.10012 | 1.5702e-06 | 4.3673e-07 | 2.3531e-07 | 1.4406e-08 | 2.2927e-08 | 9 | 2 | 0.2 | 10 | 10.012 | 0.31367 | 0.087241 |  | 0.0028778 | 0.0045799 | 0.00023503
178 |  |  | 0.10026 | 1.5947e-06 | 4.5857e-07 | 2.0817e-07 | 2.2373e-08 | 1.021e-08 | 9 | 2 | 0.2 | 10 | 10.0255 | 0.31812 | 0.091481 |  | 0.0044631 | 0.0020368 | 0.00020764
179 |  |  | 0.049526 | 8.1623e-07 | 2.2019e-07 | 1.1403e-07 | 3.0182e-08 | 1.7579e-08 | 8 | 2 | 0.2 | 10 | 4.9526 | 0.32962 | 0.088922 |  | 0.012188 | 0.007099 | 0.00023024
180 |  |  | 0.049583 | 8.418e-07 | 2.2014e-07 | 8.211e-08 | 1.2117e-08 | 1.2881e-08 | 8 | 2 | 0.2 | 10 | 4.9583 | 0.33955 | 0.088795 |  | 0.0048878 | 0.0051958 | 0.0001656
181 |  |  | 0.049637 | 8.3458e-07 | 2.2353e-07 | 9.1271e-08 | 1.0586e-08 | 2.7465e-08 | 8 | 2 | 0.2 | 10 | 4.9637 | 0.33627 | 0.090066 |  | 0.0042652 | 0.011066 | 0.00018388
182 |  |  | 0.049701 | 8.3236e-07 | 2.2559e-07 | 6.5369e-08 | 2.1928e-08 | 1.6589e-08 | 8 | 2 | 0.2 | 10 | 4.9701 | 0.33495 | 0.09078 |  | 0.0088241 | 0.0066756 | 0.00013153
183 |  |  | 0.0099302 | 1.6444e-07 | 3.1415e-08 | 1.6476e-08 | 2.8133e-08 | 1.4459e-08 | 7 | 2 | 0.2 | 10 | 0.99302 | 0.33119 | 0.063272 |  | 0.056661 | 0.029121 | 0.00016592
184 |  |  | 0.0099597 | 1.8339e-07 | 5.1054e-08 | 8.4555e-09 | 1.12e-08 | 2.0266e-08 | 7 | 2 | 0.2 | 10 | 0.99597 | 0.36825 | 0.10252 |  | 0.022491 | 0.040695 | 8.4897e-05
185 |  |  | 0.0099462 | 1.6901e-07 | 4.5721e-08 | 5.3685e-08 | 1.7729e-08 | 1.8592e-08 | 7 | 2 | 0.2 | 10 | 0.99462 | 0.33984 | 0.091937 |  | 0.035649 | 0.037385 | 0.00053975
186 |  |  | 0.0099598 | 1.7124e-07 | 5.1677e-08 | 1.0849e-07 | 1.5564e-08 | 2.6984e-08 | 7 | 2 | 0.2 | 10 | 0.99598 | 0.34386 | 0.10377 |  | 0.031254 | 0.054185 | 0.0010893
187 |  |  | 0.0099804 | 1.7536e-07 | 7.1133e-08 | 2.3139e-08 | 1.5976e-08 | 7.6006e-09 | 7 | 2 | 0.2 | 10 | 0.99804 | 0.3514 | 0.14255 |  | 0.032015 | 0.015231 | 0.00023184
188 |  |  | 0.0099784 | 1.893e-07 | 5.3601e-08 | 1.0915e-07 | 1.2087e-08 | 2.144e-08 | 7 | 2 | 0.2 | 10 | 0.99784 | 0.37942 | 0.10743 |  | 0.024226 | 0.042973 | 0.0010939
189 |  |  | 0.0099665 | 1.6992e-07 | 3.8554e-08 | 1.9041e-08 | 1.3714e-08 | 1.5187e-08 | 7 | 2 | 0.2 | 10 | 0.99665 | 0.34099 | 0.077367 |  | 0.02752 | 0.030475 | 0.00019105
190 |  |  | 0.0099796 | 1.952e-07 | 3.7418e-08 | 4.7984e-08 | 1.6108e-08 | 2.4461e-08 | 7 | 2 | 0.2 | 10 | 0.99796 | 0.39119 | 0.074988 |  | 0.032282 | 0.049022 | 0.00048082
191 |  |  | 0.0099909 | 1.7558e-07 | 4.5853e-08 | 4.0171e-08 | 1.5731e-08 | 1.2549e-08 | 7 | 2 | 0.2 | 10 | 0.99909 | 0.35148 | 0.09179 |  | 0.031491 | 0.025122 | 0.00040207
192 |  |  | 0.010009 | 1.8118e-07 | 5.1159e-08 | 2.5058e-08 | 1.4036e-08 | 1.9552e-08 | 7 | 2 | 0.2 | 10 | 1.0009 | 0.36206 | 0.10223 |  | 0.028049 | 0.039071 | 0.00025037
193 |  |  | 0.01 | 1.6593e-07 | 5.3259e-08 | 6.5266e-08 | 1.7235e-08 | 1.7069e-08 | 7 | 2 | 0.2 | 10 | 1 | 0.33185 | 0.10651 |  | 0.034468 | 0.034137 | 0.00065263
194 |  |  | 0.010004 | 1.4426e-07 | 5.5648e-08 | 5.4178e-08 | 1.8878e-08 | 1.3661e-08 | 7 | 2 | 0.2 | 10 | 1.0004 | 0.2884 | 0.11125 |  | 0.037739 | 0.02731 | 0.00054155
195
196
197 |  | Calibration of 100:5 ratio, outer primary of Ta, vs 100:5 ratio, inner primary of Tb | [8 empty] | Wagner at Tb end of link, phase SR830 to the MSL shunt.
198 |  |  | 0.10048 | 1.818e-06 | 5.4349e-07 | 3.8006e-08 | 1.9295e-08 | 1.0962e-08 | 9 | 2 | 0.2 | 10 | 10.0476 | 0.36188 | 0.10818 |  | 0.0038407 | 0.0021819 | 3.7826e-05 |  |  | Initial Setup
199 |  |  | 0.10041 | -0.00018378 | -1.2396e-05 | 5.9561e-08 | 7.3274e-08 | 2.6752e-08 | 15 | 2 | 0.2 | 10 | 10.0408 | -36.6067 | -2.4691 |  | 0.014595 | 0.0053287 | 5.9319e-05 |  |  | 10k ohm across Ta main secondary
200 |  |  | 0.10018 | 4.4918e-05 | -0.00064255 | 5.2165e-08 | 3.9534e-08 | 1.6112e-07 | 17 | 2 | 0.2 | 10 | 10.0184 | 8.9672 | -128.2735 |  | 0.0078923 | 0.032164 | 5.2069e-05 |  |  | 1uF across Ta main secondary
201 |  |  | 0.10036 | 1.8566e-06 | 5.4008e-07 | 2.5815e-08 | 3.9126e-08 | 3.8151e-08 | 10 | 2 | 0.2 | 10 | 10.0363 | 0.36998 | 0.10762 |  | 0.0077969 | 0.0076026 | 2.5722e-05
202 |  |  | 1.2511 | 1.1003e-05 | 5.4686e-06 | 9.6464e-06 | 1.4457e-07 | 2.5662e-08 | 12 | 2 | 0.2 | 10 | 125.1098 | 0.1759 | 0.08742 |  | 0.0023111 | 0.00041022 | 0.00077103 |  | Zetka clipping (Red) | Start at 125%. 3458A on 12V range.
203 |  |  | 1.2454 | 9.9078e-06 | 5.2847e-06 | 7.888e-06 | 4.9521e-08 | 1.6837e-08 | 12 | 2 | 0.2 | 10 | 124.5407 | 0.15911 | 0.084867 |  | 0.00079526 | 0.00027039 | 0.00063337 |  | Zetka clipping (Red) | 125%. 3458A on 12V range.
204 |  |  | 1.2519 | 1.0453e-05 | 5.3777e-06 | 8.5265e-06 | 7.5616e-08 | 1.195e-08 | 12 | 2 | 0.2 | 10 | 125.1915 | 0.16698 | 0.085911 |  | 0.001208 | 0.00019091 | 0.00068108 |  | Zetka clipping (Red) | 125%. 3458A on 12V range.
205 |  |  | 1.2476 | 9.8739e-06 | 5.3358e-06 | 5.5801e-06 | 7.0706e-08 | 2.9047e-08 | 12 | 2 | 0.2 | 10 | 124.7616 | 0.15828 | 0.085535 |  | 0.0011335 | 0.00046564 | 0.00044726 |  | Zetka clipping (Red) | 125%. 3458A on 12V range.
206 |  |  | 1.251 | 1.0532e-05 | 5.3385e-06 | 6.7658e-06 | 1.3905e-07 | 2.5905e-08 | 12 | 2 | 0.2 | 10 | 125.0967 | 0.16838 | 0.08535 |  | 0.002223 | 0.00041415 | 0.00054085 |  | Zetka clipping (Red) | 125%. 3458A on 12V range.
207 |  |  | 1.2478 | 1.0549e-05 | 5.3596e-06 | 5.2581e-06 | 4.1056e-08 | 1.7252e-08 | 12 | 2 | 0.2 | 10 | 124.7754 | 0.16909 | 0.085908 |  | 0.00065808 | 0.00027653 | 0.0004214 |  | Zetka clipping (Red) | 125%. 3458A on 12V range.
208 |  |  | 1.1989 | 1.0605e-05 | 5.1453e-06 | 2.457e-06 | 4.2808e-08 | 2.4759e-08 | 12 | 2 | 0.2 | 10 | 119.8887 | 0.17691 | 0.085835 |  | 0.00071413 | 0.00041304 | 0.00020494 |  | Zetka clipping (Red) | 120%. 3458A at 1.2V range
209 |  |  | 1.1981 | 1.0105e-05 | 5.1597e-06 | 5.224e-06 | 3.1554e-08 | 1.853e-08 | 12 | 2 | 0.2 | 10 | 119.8126 | 0.16868 | 0.086129 |  | 0.00052673 | 0.00030932 | 0.00043601 |  | Zetka clipping (Red)
210 |  |  | 1.1984 | 1.0371e-05 | 5.1429e-06 | 3.3107e-06 | 2.222e-08 | 3.0915e-08 | 12 | 2 | 0.2 | 10 | 119.8402 | 0.17309 | 0.085829 |  | 0.00037082 | 0.00051593 | 0.00027626 |  | Zetka clipping (Red)
211 |  |  | 1.1983 | 1.0355e-05 | 5.1005e-06 | 3.4907e-06 | 5.8563e-08 | 1.2974e-08 | 12 | 2 | 0.2 | 10 | 119.8347 | 0.17282 | 0.085125 |  | 0.0009774 | 0.00021654 | 0.0002913 |  | Zetka clipping (Red)
212 |  |  | 1.1983 | 1.0268e-05 | 5.08e-06 | 3.6708e-06 | 3.3627e-08 | 2.0184e-08 | 12 | 2 | 0.2 | 10 | 119.8292 | 0.17137 | 0.084787 |  | 0.00056124 | 0.00033688 | 0.00030633 |  | Zetka clipping (Red)
213 |  |  | 1.1993 | 1.0219e-05 | 5.0912e-06 | 3.4461e-06 | 4.0712e-08 | 1.7038e-08 | 12 | 2 | 0.2 | 10 | 119.9347 | 0.17042 | 0.084899 |  | 0.0006789 | 0.00028411 | 0.00028733 |  | Zetka clipping (Red)
214 |  |  | 1.0031 | 9.235e-06 | 4.3609e-06 | 1.8477e-07 | 5.0628e-08 | 1.7408e-08 | 11 | 2 | 0.2 | 10 | 100.3115 | 0.18413 | 0.086947 |  | 0.0010094 | 0.00034707 | 1.8419e-05 |  | Zetka clipping (Red)
215 |  |  | 1.0007 | 9.0954e-06 | 4.3312e-06 | 6.3759e-07 | 2.323e-08 | 1.8921e-08 | 11 | 2 | 0.2 | 10 | 100.0712 | 0.18178 | 0.086562 |  | 0.00046427 | 0.00037815 | 6.3714e-05 |  | Zetka clipping (Red)
216 |  |  | 0.99962 | 8.9489e-06 | 4.3577e-06 | 1.0944e-07 | 1.7271e-08 | 1.7378e-08 | 11 | 2 | 0.2 | 10 | 99.9622 | 0.17905 | 0.087187 |  | 0.00034555 | 0.00034769 | 1.0948e-05 |  | Zetka clipping (Red)
217 |  |  | 0.99957 | 8.9792e-06 | 4.3317e-06 | 1.5369e-07 | 1.9395e-08 | 1.4519e-08 | 11 | 2 | 0.2 | 10 | 99.9572 | 0.17966 | 0.086672 |  | 0.00038807 | 0.00029051 | 1.5376e-05 |  | Zetka clipping (Red)
218 |  |  | 0.60007 | 6.4302e-06 | 2.8793e-06 | 1.9729e-06 | 3.1002e-08 | 3.1709e-08 | 11 | 2 | 0.2 | 10 | 60.0071 | 0.21432 | 0.095965 |  | 0.0010333 | 0.0010568 | 0.00032877
219 |  |  | 0.60103 | 6.412e-06 | 2.8629e-06 | 1.854e-06 | 2.1585e-08 | 1.5961e-08 | 11 | 2 | 0.2 | 10 | 60.1028 | 0.21337 | 0.095267 |  | 0.00071828 | 0.00053114 | 0.00030848
220 |  |  | 0.40001 | 4.8893e-06 | 1.8202e-06 | 1.4058e-06 | 4.2981e-08 | 4.6011e-08 | 10 | 2 | 0.2 | 10 | 40.001 | 0.24446 | 0.091006 |  | 0.002149 | 0.0023005 | 0.00035145
221 |  |  | 0.401 | 4.9159e-06 | 1.7805e-06 | 1.3188e-06 | 3.3026e-08 | 3.1166e-08 | 10 | 2 | 0.2 | 10 | 40.1001 | 0.24518 | 0.088802 |  | 0.0016472 | 0.0015544 | 0.00032888
222 |  |  | 0.20003 | 2.712e-06 | 8.8067e-07 | 5.7997e-07 | 1.3119e-07 | 1.1284e-08 | 10 | 2 | 0.2 | 10 | 20.0026 | 0.27116 | 0.088056 |  | 0.013117 | 0.0011283 | 0.00028995
223 |  |  | 0.20035 | 2.808e-06 | 8.707e-07 | 5.6712e-07 | 1.4689e-08 | 1.8685e-08 | 10 | 2 | 0.2 | 10 | 20.0355 | 0.2803 | 0.086916 |  | 0.0014663 | 0.0018652 | 0.00028306
224 |  |  | 0.099827 | 1.4838e-06 | 4.3255e-07 | 2.7169e-07 | 2.1753e-08 | 2.5626e-08 | 9 | 2 | 0.2 | 10 | 9.9827 | 0.29728 | 0.08666 |  | 0.0043581 | 0.0051341 | 0.00027217
225 |  |  | 0.10001 | 1.5329e-06 | 4.0815e-07 | 2.4875e-07 | 4.2584e-08 | 3.7312e-08 | 9 | 2 | 0.2 | 10 | 10.0006 | 0.30656 | 0.081626 |  | 0.0085164 | 0.007462 | 0.00024873
226 |  |  | 0.10014 | 1.5289e-06 | 4.4814e-07 | 2.3734e-07 | 4.4436e-08 | 3.6266e-08 | 9 | 2 | 0.2 | 10 | 10.0143 | 0.30534 | 0.0895 |  | 0.0088746 | 0.0072428 | 0.000237
227 |  |  | 0.10027 | 1.5168e-06 | 4.3497e-07 | 2.2701e-07 | 5.1275e-08 | 3.7126e-08 | 9 | 2 | 0.2 | 10 | 10.0271 | 0.30255 | 0.086759 |  | 0.010227 | 0.0074052 | 0.0002264
228 |  |  | 0.04978 | 7.8448e-07 | 2.919e-07 | 1.1146e-07 | 5.9904e-08 | 4.4901e-08 | 8 | 2 | 0.2 | 10 | 4.978 | 0.31518 | 0.11727 |  | 0.024068 | 0.01804 | 0.0002239
229 |  |  | 0.04984 | 7.9108e-07 | 2.3043e-07 | 9.0731e-08 | 3.65e-08 | 3.1016e-08 | 8 | 2 | 0.2 | 10 | 4.984 | 0.31745 | 0.092465 |  | 0.014647 | 0.012446 | 0.00018204
230 |  |  | 0.049897 | 7.7993e-07 | 2.2083e-07 | 7.6314e-08 | 3.174e-08 | 2.1137e-08 | 8 | 2 | 0.2 | 10 | 4.9897 | 0.31262 | 0.088514 |  | 0.012722 | 0.0084722 | 0.00015294
231 |  |  | 0.049937 | 8.0825e-07 | 2.4218e-07 | 8.7061e-08 | 3.4121e-08 | 3.8916e-08 | 8 | 2 | 0.2 | 10 | 4.9937 | 0.32371 | 0.096996 |  | 0.013666 | 0.015586 | 0.00017434
232 |  |  | 0.049982 | 8.05e-07 | 2.2917e-07 | 8.4661e-08 | 2.3758e-08 | 2.5292e-08 | 8 | 2 | 0.2 | 10 | 4.9982 | 0.32212 | 0.091701 |  | 0.0095066 | 0.01012 | 0.00016938
233 |  |  | 0.050038 | 7.9752e-07 | 2.1148e-07 | 9.8314e-08 | 6.0311e-08 | 6.3181e-08 | 8 | 2 | 0.2 | 10 | 5.0038 | 0.31877 | 0.084529 |  | 0.024106 | 0.025253 | 0.00019648
234 |  |  | 0.050085 | 7.9775e-07 | 2.3014e-07 | 6.4303e-08 | 1.9504e-08 | 2.0406e-08 | 8 | 2 | 0.2 | 10 | 5.0085 | 0.31856 | 0.0919 |  | 0.0077884 | 0.0081484 | 0.00012839
235 |  |  | 0.050132 | 8.1712e-07 | 2.6529e-07 | 7.1805e-08 | 5.7333e-08 | 3.4141e-08 | 8 | 2 | 0.2 | 10 | 5.0132 | 0.32599 | 0.10583 |  | 0.022873 | 0.013621 | 0.00014323
236 |  |  | 0.0098821 | 1.6484e-07 | 6.3712e-08 | 5.8857e-08 | 2.8181e-08 | 3.5201e-08 | 8 | 2 | 0.2 | 10 | 0.98821 | 0.33361 | 0.12894 |  | 0.057035 | 0.071242 | 0.00059559
237 |  |  | 0.0098945 | 1.7349e-07 | 8.2059e-08 | 5.3783e-08 | 5.9372e-08 | 5.329e-08 | 8 | 2 | 0.2 | 10 | 0.98945 | 0.35068 | 0.16587 |  | 0.12001 | 0.10772 | 0.00054357
238 |  |  | 0.0098965 | 1.448e-07 | 5.4663e-08 | 1.3185e-07 | 5.7461e-08 | 4.9369e-08 | 8 | 2 | 0.2 | 10 | 0.98965 | 0.29264 | 0.11047 |  | 0.11612 | 0.099771 | 0.0013323
239 |  |  | 0.0099018 | 1.5774e-07 | 3.3435e-08 | 1.0682e-07 | 2.5657e-08 | 2.3045e-08 | 8 | 2 | 0.2 | 10 | 0.99018 | 0.3186 | 0.067533 |  | 0.051824 | 0.046547 | 0.0010788
240 |  |  | 0.0099011 | 1.6895e-07 | 4.5034e-08 | 5.4062e-08 | 5.202e-08 | 8.9646e-08 | 7 | 2 | 0.2 | 10 | 0.99011 | 0.34128 | 0.090968 |  | 0.10508 | 0.18108 | 0.00054602
241 |  |  | 0.0099122 | 1.6332e-07 | 3.9138e-08 | 7.6268e-08 | 3.0834e-08 | 4.608e-08 | 7 | 2 | 0.2 | 10 | 0.99122 | 0.32953 | 0.078968 |  | 0.062215 | 0.092976 | 0.00076943
242 |  |  | 0.009913 | 2.1119e-07 | 4.1828e-08 | 3.5958e-08 | 6.4859e-08 | 5.0372e-08 | 7 | 2 | 0.2 | 10 | 0.9913 | 0.42609 | 0.084391 |  | 0.13086 | 0.10163 | 0.00036274
243 |  |  | 0.009924 | 1.5155e-07 | 4.2426e-08 | 5.3685e-08 | 2.4984e-08 | 3.947e-08 | 7 | 2 | 0.2 | 10 | 0.9924 | 0.30542 | 0.085502 |  | 0.050349 | 0.079544 | 0.00054096
244 |  |  | 0.0099336 | 1.7047e-07 | 3.2585e-08 | 1.2119e-07 | 3.785e-08 | 3.5873e-08 | 7 | 2 | 0.2 | 10 | 0.99336 | 0.34321 | 0.065605 |  | 0.076206 | 0.072225 | 0.00122
245 |  |  | 0.0099263 | 1.3974e-07 | 4.432e-08 | 1.42e-08 | 4.7945e-08 | 5.1043e-08 | 7 | 2 | 0.2 | 10 | 0.99263 | 0.28155 | 0.089298 |  | 0.096603 | 0.10284 | 0.00014306
246 |  |  | 0.0099462 | 1.9075e-07 | 5.7737e-08 | 8.4392e-08 | 3.2582e-08 | 3.197e-08 | 7 | 2 | 0.2 | 10 | 0.99462 | 0.38356 | 0.1161 |  | 0.065515 | 0.064286 | 0.00084848
247 |  |  | 0.0099544 | 2.6193e-07 | 4.6847e-08 | 1.1298e-07 | 4.7939e-08 | 5.4728e-08 | 7 | 2 | 0.2 | 10 | 0.99544 | 0.52627 | 0.094124 |  | 0.096317 | 0.10996 | 0.001135
248 |  |  | 0.0099555 | 1.6018e-07 | 3.3644e-08 | 1.9987e-08 | 3.8751e-08 | 4.1042e-08 | 7 | 2 | 0.2 | 10 | 0.99555 | 0.32179 | 0.067589 |  | 0.077849 | 0.082451 | 0.00020077
249 |  |  | 0.0099868 | 2.2159e-07 | 8.4291e-08 | 3.7904e-08 | 7.8865e-08 | 4.8437e-08 | 7 | 2 | 0.2 | 10 | 0.99868 | 0.44377 | 0.1688 |  | 0.15794 | 0.097002 | 0.00037954
250
251
252
253 |  | Average of below | 1.2513 | -9.9664e-07 | -1.0917e-06 | 1.0376e-05 | 6.3058e-08 | 1.8242e-08 | 9 | 2 | 0.2 | 10 | 125.1251 | -0.01593 | -0.017449 | 0 | 0.0010079 | 0.00029158 | 0.00082922
254 |  |  | 1.198 | -1.1732e-06 | -1.0612e-06 | 5.7854e-06 | 1.803e-08 | 1.925e-08 | 9 | 2 | 0.2 | 10 | 119.798 | -0.019587 | -0.017716 | 0 | 0.00030101 | 0.00032138 | 0.00048293
255 |  |  | 1.0011 | -1.5706e-06 | -8.8567e-07 | 6.3748e-07 | 1.7877e-08 | 1.3466e-08 | 9 | 2 | 0.2 | 10 | 100.1098 | -0.031378 | -0.017694 | 0 | 0.00035714 | 0.00026903 | 6.3678e-05
256 |  |  | 0.60017 | -1.4923e-06 | -4.2484e-07 | 1.0631e-06 | 1.0823e-08 | 2.6198e-08 | 9 | 2 | 0.2 | 10 | 60.0168 | -0.049731 | -0.014157 | 0 | 0.00036066 | 0.00087303 | 0.00017713
257 |  |  | 0.40012 | -1.1169e-06 | -3.6244e-07 | 7.2048e-07 | 9.8564e-09 | 1.8892e-08 | 9 | 2 | 0.2 | 10 | 40.0116 | -0.055829 | -0.018117 | 0 | 0.00049268 | 0.00094432 | 0.00018007
258 |  |  | 0.19997 | -6.7534e-07 | -2.2279e-07 | 4.561e-07 | 1.5968e-08 | 1.7202e-08 | 8 | 2 | 0.2 | 10 | 19.9968 | -0.067545 | -0.022282 | 0 | 0.001597 | 0.0017204 | 0.00022809
259 |  |  | 0.10025 | -3.6929e-07 | -1.1945e-07 | 1.6843e-07 | 1.9352e-08 | 1.4687e-08 | 8 | 2 | 0.2 | 10 | 10.0245 | -0.073679 | -0.023821 | 0 | 0.0038604 | 0.0029322 | 0.00016802
260 |  |  | 0.049913 | -1.9423e-07 | -7.1304e-08 | 6.365e-08 | 4.2332e-08 | 3.6751e-08 | 7.8 | 2 | 0.2 | 10 | 4.9913 | -0.07783 | -0.028569 | 0 | 0.016951 | 0.014717 | 0.00012753
261 |  |  | 0.01001 | -4.1957e-08 | -1.0472e-08 | 5.9893e-08 | 2.6542e-08 | 2.9053e-08 | 7 | 2 | 0.2 | 10 | 1.001 | -0.083794 | -0.020985 | 0 | 0.052979 | 0.05803 | 0.00059799
262
263 |  | Copy from above | 0.10037 | -3.7501e-07 | -1.3897e-07 | 1.9686e-07 | 1.39e-08 | 1.4283e-08 | 8 | 2 | 0.2 | 10 | 10.0371 | -0.074725 | -0.027692 | 0 | 0.0027698 | 0.002846 | 0.00019613
264 |  |  | 0.10052 | -3.6226e-07 | -1.2849e-07 | 1.451e-07 | 2.4981e-08 | 1.0552e-08 | 8 | 2 | 0.2 | 10 | 10.0518 | -0.072079 | -0.025566 | 0 | 0.0049705 | 0.0020995 | 0.00014435
265 |  |  | 0.099847 | -3.706e-07 | -9.0883e-08 | 1.6334e-07 | 1.9176e-08 | 1.9226e-08 | 8 | 2 | 0.2 | 10 | 9.9847 | -0.074234 | -0.018204 | 0 | 0.003841 | 0.0038511 | 0.00016359
266 |  |  | 0.049833 | -2.0304e-07 | -8.6512e-08 | 6.5881e-08 | 2.1184e-08 | 2.1883e-08 | 8 | 2 | 0.2 | 10 | 4.9833 | -0.081487 | -0.034721 | 0 | 0.0085021 | 0.0087826 | 0.0001322
267 |  |  | 0.049867 | -1.9732e-07 | -5.1665e-08 | 6.7288e-08 | 2.5941e-08 | 2.2409e-08 | 8 | 2 | 0.2 | 10 | 4.9867 | -0.079139 | -0.020721 | 0 | 0.010404 | 0.0089873 | 0.00013493
268 |  |  | 0.0499 | -1.9916e-07 | -5.4081e-08 | 6.4394e-08 | 1.7959e-08 | 1.7726e-08 | 8 | 2 | 0.2 | 10 | 4.99 | -0.079826 | -0.021676 | 0 | 0.0071982 | 0.0071046 | 0.00012905
269 |  |  | 0.049958 | -1.7751e-07 | -7.9006e-08 | 6.7035e-08 | 7.5659e-08 | 5.5737e-08 | 8 | 2 | 0.2 | 10 | 4.9958 | -0.071064 | -0.031629 | 0 | 0.030289 | 0.022314 | 0.00013418
270 |  |  | 0.050006 | -1.9411e-07 | -8.5256e-08 | 5.3652e-08 | 7.0916e-08 | 6.5999e-08 | 7 | 2 | 0.2 | 10 | 5.0006 | -0.077635 | -0.034098 | 0 | 0.028363 | 0.026396 | 0.00010729
271 |  |  | 0.010013 | -3.8433e-08 | -2.8534e-08 | 7.3394e-08 | 2.1625e-08 | 3.7432e-08 | 7 | 2 | 0.2 | 10 | 1.0013 | -0.07677 | -0.056995 | 0 | 0.043195 | 0.07477 | 0.00073302
272 |  |  | 0.010034 | -2.6908e-08 | 2.8958e-09 | 2.8792e-08 | 2.7437e-08 | 2.5167e-08 | 7 | 2 | 0.2 | 10 | 1.0034 | -0.053636 | 0.0057723 | 0 | 0.05469 | 0.050165 | 0.00028696
273 |  |  | 0.010022 | -5.7745e-08 | -1.7865e-08 | 1.012e-07 | 2.9512e-08 | 2.9139e-08 | 7 | 2 | 0.2 | 10 | 1.0022 | -0.11524 | -0.035652 | 0 | 0.058894 | 0.05815 | 0.0010097
274 |  |  | 0.01004 | -5.3751e-08 | -7.0548e-09 | 8.8608e-08 | 4.7204e-08 | 3.3258e-08 | 7 | 2 | 0.2 | 10 | 1.004 | -0.10707 | -0.014053 | 0 | 0.094032 | 0.06625 | 0.00088254
275 |  |  | 0.010031 | -4.031e-08 | 5.6316e-09 | 5.7969e-08 | 2.406e-08 | 2.0802e-08 | 7 | 2 | 0.2 | 10 | 1.0031 | -0.080371 | 0.011228 | 0 | 0.047972 | 0.041475 | 0.0005779
276 |  |  | 0.010045 | -5.5979e-08 | -8.5333e-09 | 4.3181e-08 | 4.7496e-08 | 5.645e-08 | 7 | 2 | 0.2 | 10 | 1.0045 | -0.11146 | -0.01699 | 0 | 0.094567 | 0.1124 | 0.00042988
277 |  |  | 0.010037 | -3.7645e-08 | -5.7291e-09 | 4.879e-08 | 2.8198e-08 | 1.9176e-08 | 7 | 2 | 0.2 | 10 | 1.0037 | -0.075013 | -0.011416 | 0 | 0.056189 | 0.038211 | 0.00048612
278 |  |  | 0.010041 | -4.4749e-08 | -9.0106e-09 | 2.0761e-08 | 1.3271e-08 | 2.2271e-08 | 7 | 2 | 0.2 | 10 | 1.0041 | -0.089133 | -0.017948 | 0 | 0.026435 | 0.044361 | 0.00020677
279 |  |  | 0.010064 | -3.6278e-08 | -5.9329e-09 | 1.0046e-07 | 1.8937e-08 | 2.564e-08 | 7 | 2 | 0.2 | 10 | 1.0064 | -0.072096 | -0.01179 | 0 | 0.037634 | 0.050955 | 0.00099826
280 |  |  | 0.0099026 | -3.9023e-08 | -2.5847e-08 | 6.7076e-08 | 2.1796e-08 | 2.2097e-08 | 7 | 2 | 0.2 | 10 | 0.99026 | -0.078813 | -0.052203 | 0 | 0.044021 | 0.044629 | 0.00067736
281 |  |  | 0.0098829 | -3.0703e-08 | -1.5213e-08 | 2.8596e-08 | 1.2422e-08 | 2.815e-08 | 7 | 2 | 0.2 | 10 | 0.98829 | -0.062134 | -0.030786 | 0 | 0.025139 | 0.056967 | 0.00028935
282
283
284 |  | Average of below | 1.2469 | 1.0782e-05 | 4.8336e-06 | 8.386e-06 | 1.3801e-07 | 6.7359e-08 | 12 | 2 | 0.2 | 10 | 124.695 | 0.17294 | 0.077527 | 0 | 0.0022154 | 0.0010822 | 0.00067254
285 |  |  | 1.1987 | 1.1e-05 | 5.0788e-06 | 5.5892e-06 | 1.1289e-07 | 1.956e-08 | 12 | 2 | 0.2 | 10 | 119.8662 | 0.18354 | 0.084741 | 0 | 0.0018835 | 0.00032636 | 0.00046631
286 |  |  | 1.0006 | 9.5729e-06 | 4.3306e-06 | 4.659e-07 | 2.2943e-08 | 1.7106e-08 | 11 | 2 | 0.2 | 10 | 100.0641 | 0.19133 | 0.086556 | 0 | 0.00045854 | 0.0003419 | 4.656e-05
287 |  |  | 0.60001 | 6.5245e-06 | 3.0061e-06 | 1.3435e-06 | 2.7138e-08 | 2.2329e-08 | 11 | 2 | 0.2 | 10 | 60.001 | 0.21748 | 0.1002 | 0 | 0.00090473 | 0.00074428 | 0.00022391
288 |  |  | 0.39993 | 5.069e-06 | 1.9178e-06 | 1.1242e-06 | 5.445e-08 | 4.8145e-08 | 10 | 2 | 0.2 | 10 | 39.9926 | 0.2535 | 0.095908 | 0 | 0.0027221 | 0.0024069 | 0.00028111
289 |  |  | 0.20005 | 2.9039e-06 | 9.3336e-07 | 5.2398e-07 | 4.5633e-08 | 3.8227e-08 | 10 | 2 | 0.2 | 10 | 20.0052 | 0.29032 | 0.093311 | 0 | 0.0045641 | 0.0038229 | 0.00026194
290 |  |  | 0.10012 | 1.5889e-06 | 4.3427e-07 | 2.3054e-07 | 2.0378e-08 | 1.835e-08 | 9 | 2 | 0.2 | 10 | 10.0116 | 0.3174 | 0.086749 | 0 | 0.004071 | 0.0036669 | 0.00023029
291 |  |  | 0.049611 | 8.3124e-07 | 2.2236e-07 | 8.8194e-08 | 1.8703e-08 | 1.8629e-08 | 8 | 2 | 0.2 | 10 | 4.9611 | 0.3351 | 0.089641 | 0 | 0.0075414 | 0.0075092 | 0.00017781
292 |  |  | 0.0099754 | 1.7373e-07 | 4.8874e-08 | 4.7591e-08 | 1.6366e-08 | 1.7652e-08 | 7 | 2 | 0.2 | 10 | 0.99754 | 0.34833 | 0.097968 | 0 | 0.03282 | 0.035394 | 0.00047701
293
294 |  | Copy from above | 1.2499 | 1.0617e-05 | 4.8248e-06 | 8.2784e-06 | 8.8289e-08 | 1.9778e-08 | 12 | 2 | 0.2 | 10 | 124.9885 | 0.16989 | 0.077204 | 0 | 0.0014128 | 0.00031647 | 0.00066233
295 |  |  | 1.244 | 1.0946e-05 | 4.8424e-06 | 8.4936e-06 | 1.8773e-07 | 1.1494e-07 | 12 | 2 | 0.2 | 10 | 124.4014 | 0.17599 | 0.077851 | 0 | 0.0030181 | 0.0018479 | 0.00068275
296 |  | anomalous? | 1.2486 | 1.8662e-05 | 5.9264e-06 | 1.1288e-05 | 2.3455e-06 | 5.4105e-07 | 12 | 2 | 0.2 | 10 | 124.8588 | 0.29894 | 0.09493 | 0 | 0.03757 | 0.0086666 | 0.00090407
297 |  | anomalous? | 1.2475 | 1.9172e-05 | 6.1008e-06 | 6.6031e-06 | 3.7974e-06 | 4.7212e-07 | 13 | 2 | 0.2 | 10 | 124.7492 | 0.30737 | 0.09781 | 0 | 0.06088 | 0.0075691 | 0.00052931
298 |  |  | 1.1988 | 1.1119e-05 | 5.1091e-06 | 4.8775e-06 | 3.3207e-07 | 2.165e-08 | 12 | 2 | 0.2 | 10 | 119.8778 | 0.1855 | 0.085238 | 0 | 0.0055402 | 0.0003612 | 0.00040687
299 |  |  | 1.1983 | 1.1127e-05 | 5.0826e-06 | 7.5549e-06 | 5.6332e-08 | 1.4931e-08 | 12 | 2 | 0.2 | 10 | 119.8308 | 0.18571 | 0.08483 | 0 | 0.00094019 | 0.00024921 | 0.00063046
300 |  |  | 1.1986 | 1.0981e-05 | 5.1312e-06 | 5.8264e-06 | 3.4197e-08 | 2.3801e-08 | 12 | 2 | 0.2 | 10 | 119.8624 | 0.18323 | 0.085617 | 0 | 0.0005706 | 0.00039714 | 0.00048609
301 |  |  | 1.1989 | 1.0774e-05 | 4.9922e-06 | 4.098e-06 | 2.8946e-08 | 1.7859e-08 | 12 | 2 | 0.2 | 10 | 119.894 | 0.17973 | 0.083277 | 0 | 0.00048286 | 0.00029791 | 0.0003418
302 |  |  | 1.0008 | 9.6497e-06 | 4.3417e-06 | 4.3895e-07 | 3.0068e-08 | 1.7186e-08 | 11 | 2 | 0.2 | 10 | 100.0794 | 0.19284 | 0.086765 | 0 | 0.00060088 | 0.00034344 | 4.386e-05
303 |  |  | 1.0005 | 9.496e-06 | 4.3195e-06 | 4.9285e-07 | 1.5818e-08 | 1.7026e-08 | 11 | 2 | 0.2 | 10 | 100.0488 | 0.18983 | 0.086348 | 0 | 0.0003162 | 0.00034036 | 4.9261e-05
304 |  |  | 0.59958 | 6.534e-06 | 3.0132e-06 | 1.3428e-06 | 3.3314e-08 | 2.1347e-08 | 11 | 2 | 0.2 | 10 | 59.9578 | 0.21795 | 0.10051 | 0 | 0.0011113 | 0.00071206 | 0.00022396
305 |  |  | 0.60044 | 6.515e-06 | 2.9991e-06 | 1.3442e-06 | 2.0962e-08 | 2.3312e-08 | 11 | 2 | 0.2 | 10 | 60.0442 | 0.21701 | 0.099896 | 0 | 0.0006982 | 0.00077649 | 0.00022387
306 |  |  | 0.39964 | 5.0653e-06 | 1.8919e-06 | 1.1568e-06 | 3.0164e-08 | 2.669e-08 | 10 | 2 | 0.2 | 10 | 39.964 | 0.25349 | 0.094681 | 0 | 0.0015096 | 0.0013357 | 0.00028945
307 |  |  | 0.40021 | 5.0728e-06 | 1.9438e-06 | 1.0917e-06 | 7.8735e-08 | 6.96e-08 | 10 | 2 | 0.2 | 10 | 40.0212 | 0.25351 | 0.097136 | 0 | 0.0039347 | 0.0034781 | 0.00027278
308 |  |  | 0.1999 | 2.8775e-06 | 9.294e-07 | 5.5785e-07 | 7.1279e-08 | 5.4318e-08 | 10 | 2 | 0.2 | 10 | 19.9899 | 0.2879 | 0.092986 | 0 | 0.0071315 | 0.0054345 | 0.00027907
309 |  |  | 0.20021 | 2.9304e-06 | 9.3732e-07 | 4.9012e-07 | 1.9988e-08 | 2.2136e-08 | 10 | 2 | 0.2 | 10 | 20.0205 | 0.29274 | 0.093636 | 0 | 0.0019967 | 0.0022114 | 0.00024481
310 |  |  | 0.099972 | 1.6016e-06 | 4.0752e-07 | 2.4813e-07 | 2.4354e-08 | 2.1913e-08 | 9 | 2 | 0.2 | 10 | 9.9972 | 0.32042 | 0.081526 | 0 | 0.0048721 | 0.0043839 | 0.0002482
311 |  |  | 0.10012 | 1.5702e-06 | 4.3673e-07 | 2.3531e-07 | 1.4406e-08 | 2.2927e-08 | 9 | 2 | 0.2 | 10 | 10.012 | 0.31367 | 0.087241 | 0 | 0.0028778 | 0.0045799 | 0.00023503
312 |  |  | 0.10026 | 1.5947e-06 | 4.5857e-07 | 2.0817e-07 | 2.2373e-08 | 1.021e-08 | 9 | 2 | 0.2 | 10 | 10.0255 | 0.31812 | 0.091481 | 0 | 0.0044631 | 0.0020368 | 0.00020764
313 |  |  | 0.049526 | 8.1623e-07 | 2.2019e-07 | 1.1403e-07 | 3.0182e-08 | 1.7579e-08 | 8 | 2 | 0.2 | 10 | 4.9526 | 0.32962 | 0.088922 | 0 | 0.012188 | 0.007099 | 0.00023024
314 |  |  | 0.049583 | 8.418e-07 | 2.2014e-07 | 8.211e-08 | 1.2117e-08 | 1.2881e-08 | 8 | 2 | 0.2 | 10 | 4.9583 | 0.33955 | 0.088795 | 0 | 0.0048878 | 0.0051958 | 0.0001656
315 |  |  | 0.049637 | 8.3458e-07 | 2.2353e-07 | 9.1271e-08 | 1.0586e-08 | 2.7465e-08 | 8 | 2 | 0.2 | 10 | 4.9637 | 0.33627 | 0.090066 | 0 | 0.0042652 | 0.011066 | 0.00018388
316 |  |  | 0.049701 | 8.3236e-07 | 2.2559e-07 | 6.5369e-08 | 2.1928e-08 | 1.6589e-08 | 8 | 2 | 0.2 | 10 | 4.9701 | 0.33495 | 0.09078 | 0 | 0.0088241 | 0.0066756 | 0.00013153
317 |  |  | 0.0099302 | 1.6444e-07 | 3.1415e-08 | 1.6476e-08 | 2.8133e-08 | 1.4459e-08 | 7 | 2 | 0.2 | 10 | 0.99302 | 0.33119 | 0.063272 | 0 | 0.056661 | 0.029121 | 0.00016592
318 |  |  | 0.0099597 | 1.8339e-07 | 5.1054e-08 | 8.4555e-09 | 1.12e-08 | 2.0266e-08 | 7 | 2 | 0.2 | 10 | 0.99597 | 0.36825 | 0.10252 | 0 | 0.022491 | 0.040695 | 8.4897e-05
319 |  |  | 0.0099462 | 1.6901e-07 | 4.5721e-08 | 5.3685e-08 | 1.7729e-08 | 1.8592e-08 | 7 | 2 | 0.2 | 10 | 0.99462 | 0.33984 | 0.091937 | 0 | 0.035649 | 0.037385 | 0.00053975
320 |  |  | 0.0099598 | 1.7124e-07 | 5.1677e-08 | 1.0849e-07 | 1.5564e-08 | 2.6984e-08 | 7 | 2 | 0.2 | 10 | 0.99598 | 0.34386 | 0.10377 | 0 | 0.031254 | 0.054185 | 0.0010893
321 |  |  | 0.0099804 | 1.7536e-07 | 7.1133e-08 | 2.3139e-08 | 1.5976e-08 | 7.6006e-09 | 7 | 2 | 0.2 | 10 | 0.99804 | 0.3514 | 0.14255 | 0 | 0.032015 | 0.015231 | 0.00023184
322 |  |  | 0.0099784 | 1.893e-07 | 5.3601e-08 | 1.0915e-07 | 1.2087e-08 | 2.144e-08 | 7 | 2 | 0.2 | 10 | 0.99784 | 0.37942 | 0.10743 | 0 | 0.024226 | 0.042973 | 0.0010939
323 |  |  | 0.0099665 | 1.6992e-07 | 3.8554e-08 | 1.9041e-08 | 1.3714e-08 | 1.5187e-08 | 7 | 2 | 0.2 | 10 | 0.99665 | 0.34099 | 0.077367 | 0 | 0.02752 | 0.030475 | 0.00019105
324 |  |  | 0.0099796 | 1.952e-07 | 3.7418e-08 | 4.7984e-08 | 1.6108e-08 | 2.4461e-08 | 7 | 2 | 0.2 | 10 | 0.99796 | 0.39119 | 0.074988 | 0 | 0.032282 | 0.049022 | 0.00048082
325 |  |  | 0.0099909 | 1.7558e-07 | 4.5853e-08 | 4.0171e-08 | 1.5731e-08 | 1.2549e-08 | 7 | 2 | 0.2 | 10 | 0.99909 | 0.35148 | 0.09179 | 0 | 0.031491 | 0.025122 | 0.00040207
326 |  |  | 0.010009 | 1.8118e-07 | 5.1159e-08 | 2.5058e-08 | 1.4036e-08 | 1.9552e-08 | 7 | 2 | 0.2 | 10 | 1.0009 | 0.36206 | 0.10223 | 0 | 0.028049 | 0.039071 | 0.00025037
327 |  |  | 0.01 | 1.6593e-07 | 5.3259e-08 | 6.5266e-08 | 1.7235e-08 | 1.7069e-08 | 7 | 2 | 0.2 | 10 | 1 | 0.33185 | 0.10651 | 0 | 0.034468 | 0.034137 | 0.00065263
328 |  |  | 0.010004 | 1.4426e-07 | 5.5648e-08 | 5.4178e-08 | 1.8878e-08 | 1.3661e-08 | 7 | 2 | 0.2 | 10 | 1.0004 | 0.2884 | 0.11125 | 0 | 0.037739 | 0.02731 | 0.00054155
329
330
331 |  | Average of below | 1.2491 | 1.0386e-05 | 5.3608e-06 | 7.2775e-06 | 8.6752e-08 | 2.1109e-08 | 12 | 2 | 0.2 | 10 | 124.9126 | 0.16629 | 0.085832 | 0 | 0.0013882 | 0.00033797 | 0.0005825
332 |  |  | 1.1986 | 1.0321e-05 | 5.1199e-06 | 3.5999e-06 | 3.8247e-08 | 2.0733e-08 | 12 | 2 | 0.2 | 10 | 119.8567 | 0.17221 | 0.085434 | 0 | 0.0006382 | 0.00034597 | 0.00030036
333 |  |  | 1.0008 | 9.0646e-06 | 4.3454e-06 | 2.7137e-07 | 2.7631e-08 | 1.7056e-08 | 11 | 2 | 0.2 | 10 | 100.0755 | 0.18115 | 0.086842 | 0 | 0.00055183 | 0.00034085 | 2.7114e-05
334 |  |  | 0.60055 | 6.4211e-06 | 2.8711e-06 | 1.9134e-06 | 2.6294e-08 | 2.3835e-08 | 11 | 2 | 0.2 | 10 | 60.0549 | 0.21384 | 0.095616 | 0 | 0.00087578 | 0.00079399 | 0.00031862
335 |  |  | 0.40051 | 4.9026e-06 | 1.8003e-06 | 1.3623e-06 | 3.8003e-08 | 3.8588e-08 | 10 | 2 | 0.2 | 10 | 40.0505 | 0.24482 | 0.089904 | 0 | 0.0018981 | 0.0019274 | 0.00034017
336 |  |  | 0.20019 | 2.76e-06 | 8.7569e-07 | 5.7354e-07 | 7.2938e-08 | 1.4985e-08 | 10 | 2 | 0.2 | 10 | 20.019 | 0.27573 | 0.087486 | 0 | 0.0072916 | 0.0014967 | 0.0002865
337 |  |  | 0.10006 | 1.5156e-06 | 4.3095e-07 | 2.462e-07 | 4.0012e-08 | 3.4083e-08 | 9 | 2 | 0.2 | 10 | 10.0062 | 0.30293 | 0.086136 | 0 | 0.0079941 | 0.006811 | 0.00024607
338 |  |  | 0.049961 | 7.9764e-07 | 2.4018e-07 | 8.5581e-08 | 4.0396e-08 | 3.4874e-08 | 8 | 2 | 0.2 | 10 | 4.9961 | 0.3193 | 0.096152 | 0 | 0.016172 | 0.013961 | 0.00017134
339 |  |  | 0.0099234 | 1.7718e-07 | 5.0123e-08 | 6.871e-08 | 4.4807e-08 | 4.6398e-08 | 7.2857 | 2 | 0.2 | 10 | 0.99234 | 0.357 | 0.10101 | 0 | 0.09028 | 0.09352 | 0.00069259
340
341 |  | Copy from above | 1.2511 | 1.1003e-05 | 5.4686e-06 | 9.6464e-06 | 1.4457e-07 | 2.5662e-08 | 12 | 2 | 0.2 | 10 | 125.1098 | 0.1759 | 0.08742 | 0 | 0.0023111 | 0.00041022 | 0.00077103
342 |  |  | 1.2454 | 9.9078e-06 | 5.2847e-06 | 7.888e-06 | 4.9521e-08 | 1.6837e-08 | 12 | 2 | 0.2 | 10 | 124.5407 | 0.15911 | 0.084867 | 0 | 0.00079526 | 0.00027039 | 0.00063337
343 |  |  | 1.2519 | 1.0453e-05 | 5.3777e-06 | 8.5265e-06 | 7.5616e-08 | 1.195e-08 | 12 | 2 | 0.2 | 10 | 125.1915 | 0.16698 | 0.085911 | 0 | 0.001208 | 0.00019091 | 0.00068108
344 |  |  | 1.2476 | 9.8739e-06 | 5.3358e-06 | 5.5801e-06 | 7.0706e-08 | 2.9047e-08 | 12 | 2 | 0.2 | 10 | 124.7616 | 0.15828 | 0.085535 | 0 | 0.0011335 | 0.00046564 | 0.00044726
345 |  |  | 1.251 | 1.0532e-05 | 5.3385e-06 | 6.7658e-06 | 1.3905e-07 | 2.5905e-08 | 12 | 2 | 0.2 | 10 | 125.0967 | 0.16838 | 0.08535 | 0 | 0.002223 | 0.00041415 | 0.00054085
346 |  |  | 1.2478 | 1.0549e-05 | 5.3596e-06 | 5.2581e-06 | 4.1056e-08 | 1.7252e-08 | 12 | 2 | 0.2 | 10 | 124.7754 | 0.16909 | 0.085908 | 0 | 0.00065808 | 0.00027653 | 0.0004214
347 |  |  | 1.1989 | 1.0605e-05 | 5.1453e-06 | 2.457e-06 | 4.2808e-08 | 2.4759e-08 | 12 | 2 | 0.2 | 10 | 119.8887 | 0.17691 | 0.085835 | 0 | 0.00071413 | 0.00041304 | 0.00020494
348 |  |  | 1.1981 | 1.0105e-05 | 5.1597e-06 | 5.224e-06 | 3.1554e-08 | 1.853e-08 | 12 | 2 | 0.2 | 10 | 119.8126 | 0.16868 | 0.086129 | 0 | 0.00052673 | 0.00030932 | 0.00043601
349 |  |  | 1.1984 | 1.0371e-05 | 5.1429e-06 | 3.3107e-06 | 2.222e-08 | 3.0915e-08 | 12 | 2 | 0.2 | 10 | 119.8402 | 0.17309 | 0.085829 | 0 | 0.00037082 | 0.00051593 | 0.00027626
350 |  |  | 1.1983 | 1.0355e-05 | 5.1005e-06 | 3.4907e-06 | 5.8563e-08 | 1.2974e-08 | 12 | 2 | 0.2 | 10 | 119.8347 | 0.17282 | 0.085125 | 0 | 0.0009774 | 0.00021654 | 0.0002913
351 |  |  | 1.1983 | 1.0268e-05 | 5.08e-06 | 3.6708e-06 | 3.3627e-08 | 2.0184e-08 | 12 | 2 | 0.2 | 10 | 119.8292 | 0.17137 | 0.084787 | 0 | 0.00056124 | 0.00033688 | 0.00030633
352 |  |  | 1.1993 | 1.0219e-05 | 5.0912e-06 | 3.4461e-06 | 4.0712e-08 | 1.7038e-08 | 12 | 2 | 0.2 | 10 | 119.9347 | 0.17042 | 0.084899 | 0 | 0.0006789 | 0.00028411 | 0.00028733
353 |  |  | 1.0031 | 9.235e-06 | 4.3609e-06 | 1.8477e-07 | 5.0628e-08 | 1.7408e-08 | 11 | 2 | 0.2 | 10 | 100.3115 | 0.18413 | 0.086947 | 0 | 0.0010094 | 0.00034707 | 1.8419e-05
354 |  |  | 1.0007 | 9.0954e-06 | 4.3312e-06 | 6.3759e-07 | 2.323e-08 | 1.8921e-08 | 11 | 2 | 0.2 | 10 | 100.0712 | 0.18178 | 0.086562 | 0 | 0.00046427 | 0.00037815 | 6.3714e-05
355 |  |  | 0.99962 | 8.9489e-06 | 4.3577e-06 | 1.0944e-07 | 1.7271e-08 | 1.7378e-08 | 11 | 2 | 0.2 | 10 | 99.9622 | 0.17905 | 0.087187 | 0 | 0.00034555 | 0.00034769 | 1.0948e-05
356 |  |  | 0.99957 | 8.9792e-06 | 4.3317e-06 | 1.5369e-07 | 1.9395e-08 | 1.4519e-08 | 11 | 2 | 0.2 | 10 | 99.9572 | 0.17966 | 0.086672 | 0 | 0.00038807 | 0.00029051 | 1.5376e-05
357 |  |  | 0.60007 | 6.4302e-06 | 2.8793e-06 | 1.9729e-06 | 3.1002e-08 | 3.1709e-08 | 11 | 2 | 0.2 | 10 | 60.0071 | 0.21432 | 0.095965 | 0 | 0.0010333 | 0.0010568 | 0.00032877
358 |  |  | 0.60103 | 6.412e-06 | 2.8629e-06 | 1.854e-06 | 2.1585e-08 | 1.5961e-08 | 11 | 2 | 0.2 | 10 | 60.1028 | 0.21337 | 0.095267 | 0 | 0.00071828 | 0.00053114 | 0.00030848
359 |  |  | 0.40001 | 4.8893e-06 | 1.8202e-06 | 1.4058e-06 | 4.2981e-08 | 4.6011e-08 | 10 | 2 | 0.2 | 10 | 40.001 | 0.24446 | 0.091006 | 0 | 0.002149 | 0.0023005 | 0.00035145
360 |  |  | 0.401 | 4.9159e-06 | 1.7805e-06 | 1.3188e-06 | 3.3026e-08 | 3.1166e-08 | 10 | 2 | 0.2 | 10 | 40.1001 | 0.24518 | 0.088802 | 0 | 0.0016472 | 0.0015544 | 0.00032888
361 |  |  | 0.20003 | 2.712e-06 | 8.8067e-07 | 5.7997e-07 | 1.3119e-07 | 1.1284e-08 | 10 | 2 | 0.2 | 10 | 20.0026 | 0.27116 | 0.088056 | 0 | 0.013117 | 0.0011283 | 0.00028995
362 |  |  | 0.20035 | 2.808e-06 | 8.707e-07 | 5.6712e-07 | 1.4689e-08 | 1.8685e-08 | 10 | 2 | 0.2 | 10 | 20.0355 | 0.2803 | 0.086916 | 0 | 0.0014663 | 0.0018652 | 0.00028306
363 |  |  | 0.099827 | 1.4838e-06 | 4.3255e-07 | 2.7169e-07 | 2.1753e-08 | 2.5626e-08 | 9 | 2 | 0.2 | 10 | 9.9827 | 0.29728 | 0.08666 | 0 | 0.0043581 | 0.0051341 | 0.00027217
364 |  |  | 0.10001 | 1.5329e-06 | 4.0815e-07 | 2.4875e-07 | 4.2584e-08 | 3.7312e-08 | 9 | 2 | 0.2 | 10 | 10.0006 | 0.30656 | 0.081626 | 0 | 0.0085164 | 0.007462 | 0.00024873
365 |  |  | 0.10014 | 1.5289e-06 | 4.4814e-07 | 2.3734e-07 | 4.4436e-08 | 3.6266e-08 | 9 | 2 | 0.2 | 10 | 10.0143 | 0.30534 | 0.0895 | 0 | 0.0088746 | 0.0072428 | 0.000237
366 |  |  | 0.10027 | 1.5168e-06 | 4.3497e-07 | 2.2701e-07 | 5.1275e-08 | 3.7126e-08 | 9 | 2 | 0.2 | 10 | 10.0271 | 0.30255 | 0.086759 | 0 | 0.010227 | 0.0074052 | 0.0002264
367 |  |  | 0.04978 | 7.8448e-07 | 2.919e-07 | 1.1146e-07 | 5.9904e-08 | 4.4901e-08 | 8 | 2 | 0.2 | 10 | 4.978 | 0.31518 | 0.11727 | 0 | 0.024068 | 0.01804 | 0.0002239
368 |  |  | 0.04984 | 7.9108e-07 | 2.3043e-07 | 9.0731e-08 | 3.65e-08 | 3.1016e-08 | 8 | 2 | 0.2 | 10 | 4.984 | 0.31745 | 0.092465 | 0 | 0.014647 | 0.012446 | 0.00018204
|  |  | 0.049897 | 7.7993e-07 | 2.2083e-07 | 7.6314e-08 | 3.174e-08 | 2.1137e-08 | 8 | 2 | 0.2 | 10 | 4.9897 | 0.31262 | 0.088514 | 0 | 0.012722 | 0.0084722 | 0.00015294
|  |  | 0.049937 | 8.0825e-07 | 2.4218e-07 | 8.7061e-08 | 3.4121e-08 | 3.8916e-08 | 8 | 2 | 0.2 | 10 | 4.9937 | 0.32371 | 0.096996 | 0 | 0.013666 | 0.015586 | 0.00017434
|  |  | 0.049982 | 8.05e-07 | 2.2917e-07 | 8.4661e-08 | 2.3758e-08 | 2.5292e-08 | 8 | 2 | 0.2 | 10 | 4.9982 | 0.32212 | 0.091701 | 0 | 0.0095066 | 0.01012 | 0.00016938
|  |  | 0.050038 | 7.9752e-07 | 2.1148e-07 | 9.8314e-08 | 6.0311e-08 | 6.3181e-08 | 8 | 2 | 0.2 | 10 | 5.0038 | 0.31877 | 0.084529 | 0 | 0.024106 | 0.025253 | 0.00019648
|  |  | 0.050085 | 7.9775e-07 | 2.3014e-07 | 6.4303e-08 | 1.9504e-08 | 2.0406e-08 | 8 | 2 | 0.2 | 10 | 5.0085 | 0.31856 | 0.0919 | 0 | 0.0077884 | 0.0081484 | 0.00012839
|  |  | 0.050132 | 8.1712e-07 | 2.6529e-07 | 7.1805e-08 | 5.7333e-08 | 3.4141e-08 | 8 | 2 | 0.2 | 10 | 5.0132 | 0.32599 | 0.10583 | 0 | 0.022873 | 0.013621 | 0.00014323
|  |  | 0.0098821 | 1.6484e-07 | 6.3712e-08 | 5.8857e-08 | 2.8181e-08 | 3.5201e-08 | 8 | 2 | 0.2 | 10 | 0.98821 | 0.33361 | 0.12894 | 0 | 0.057035 | 0.071242 | 0.00059559
|  |  | 0.0098945 | 1.7349e-07 | 8.2059e-08 | 5.3783e-08 | 5.9372e-08 | 5.329e-08 | 8 | 2 | 0.2 | 10 | 0.98945 | 0.35068 | 0.16587 | 0 | 0.12001 | 0.10772 | 0.00054357
|  |  | 0.0098965 | 1.448e-07 | 5.4663e-08 | 1.3185e-07 | 5.7461e-08 | 4.9369e-08 | 8 | 2 | 0.2 | 10 | 0.98965 | 0.29264 | 0.11047 | 0 | 0.11612 | 0.099771 | 0.0013323
|  |  | 0.0099018 | 1.5774e-07 | 3.3435e-08 | 1.0682e-07 | 2.5657e-08 | 2.3045e-08 | 8 | 2 | 0.2 | 10 | 0.99018 | 0.3186 | 0.067533 | 0 | 0.051824 | 0.046547 | 0.0010788
|  |  | 0.0099011 | 1.6895e-07 | 4.5034e-08 | 5.4062e-08 | 5.202e-08 | 8.9646e-08 | 7 | 2 | 0.2 | 10 | 0.99011 | 0.34128 | 0.090968 | 0 | 0.10508 | 0.18108 | 0.00054602
|  |  | 0.0099122 | 1.6332e-07 | 3.9138e-08 | 7.6268e-08 | 3.0834e-08 | 4.608e-08 | 7 | 2 | 0.2 | 10 | 0.99122 | 0.32953 | 0.078968 | 0 | 0.062215 | 0.092976 | 0.00076943
|  |  | 0.009913 | 2.1119e-07 | 4.1828e-08 | 3.5958e-08 | 6.4859e-08 | 5.0372e-08 | 7 | 2 | 0.2 | 10 | 0.9913 | 0.42609 | 0.084391 | 0 | 0.13086 | 0.10163 | 0.00036274
|  |  | 0.009924 | 1.5155e-07 | 4.2426e-08 | 5.3685e-08 | 2.4984e-08 | 3.947e-08 | 7 | 2 | 0.2 | 10 | 0.9924 | 0.30542 | 0.085502 | 0 | 0.050349 | 0.079544 | 0.00054096
|  |  | 0.0099336 | 1.7047e-07 | 3.2585e-08 | 1.2119e-07 | 3.785e-08 | 3.5873e-08 | 7 | 2 | 0.2 | 10 | 0.99336 | 0.34321 | 0.065605 | 0 | 0.076206 | 0.072225 | 0.00122
|  |  | 0.0099263 | 1.3974e-07 | 4.432e-08 | 1.42e-08 | 4.7945e-08 | 5.1043e-08 | 7 | 2 | 0.2 | 10 | 0.99263 | 0.28155 | 0.089298 | 0 | 0.096603 | 0.10284 | 0.00014306
|  |  | 0.0099462 | 1.9075e-07 | 5.7737e-08 | 8.4392e-08 | 3.2582e-08 | 3.197e-08 | 7 | 2 | 0.2 | 10 | 0.99462 | 0.38356 | 0.1161 | 0 | 0.065515 | 0.064286 | 0.00084848
|  |  | 0.0099544 | 2.6193e-07 | 4.6847e-08 | 1.1298e-07 | 4.7939e-08 | 5.4728e-08 | 7 | 2 | 0.2 | 10 | 0.99544 | 0.52627 | 0.094124 | 0 | 0.096317 | 0.10996 | 0.001135
|  |  | 0.0099555 | 1.6018e-07 | 3.3644e-08 | 1.9987e-08 | 3.8751e-08 | 4.1042e-08 | 7 | 2 | 0.2 | 10 | 0.99555 | 0.32179 | 0.067589 | 0 | 0.077849 | 0.082451 | 0.00020077
|  |  | 0.0099868 | 2.2159e-07 | 8.4291e-08 | 3.7904e-08 | 7.8865e-08 | 4.8437e-08 | 7 | 2 | 0.2 | 10 | 0.99868 | 0.44377 | 0.1688 | 0 | 0.15794 | 0.097002 | 0.00037954

## Data(2018)
Ta 5:5 SELF CALIBRATION AT 53 Hz
Index | Date | Conditions, comments | Shunt | X | Y | Sh V | X stdev | Y stdev | SR830 | SR830 | [6 empty] | SD ppm | SD ppm | SD ppm | noise sd | noise sd
|  |  | Volts | V | V | stdev/100 | V | V | Gain | Reserve | Shunt | Com R | % load | X ppm | Y ppm |  | X | Y | shunt V | X | Y |  |  |  |  | 7 Feb results |  |  | This Cal - Last Cal
1 | 2018-06-27 | 53 Hz zero burden on aux sec | 0.10061 | -2.8022e-06 | -9.6084e-07 | 2.1438e-08 | 2.4579e-08 | 5.2931e-08 | 10 | 1 | 0.2 | 10 | 10.0612 | -0.55703 | -0.191 |  | 0.004886 | 0.010522 | 2.1308e-05 | 0.019878 | 0.019878 |  | Initial Setup |  |  | -0.20844 | -0.1854 |  | -0.34859 | -0.0056026
2 |  |  | 0.10061 | 0.00011833 | 4.0971e-06 | 1.9564e-08 | 3.8588e-09 | 1.1902e-07 | 15 | 1 | 0.2 | 10 | 10.0608 | 23.5222 | 0.81447 |  | 0.0007671 | 0.023661 | 1.9445e-05 | 0.019879 | 0.019879 |  | 1k ohm across main secondary |  |  | -0.2352 | -0.19115 |  | 23.7574 | 1.0056
3 |  |  | 0.10059 | -1.4731e-05 | 0.00039861 | 1.3192e-08 | 4.2787e-07 | 3.5237e-18 | 16 | 1 | 0.2 | 10 | 10.0594 | -2.9288 | 79.2508 |  | 0.085068 | 7.0057e-13 | 1.3114e-05 | 0.019882 | 0.019882 |  | 1uF across main secondary |  |  | -0.30969 | -0.17776 |  | -2.6191 | 79.4286
4 |  |  | 1.1957 | -4.3476e-05 | -1.7089e-05 | 9.0433e-06 | 3.6259e-06 | 6.298e-07 | 13 | 1 | 0.2 | 10 | 119.5664 | -0.72723 | -0.28584 |  | 0.06065 | 0.010535 | 0.00075634 | 0.0016727 | 0.0016727 |  | Start at 120%. 3458A on 1V range. |  |  | -0.36413 | -0.17167 |  | -0.10238 | -0.082828
5 |  |  | 1.0028 | -2.3391e-05 | -1.2761e-05 | 1.0142e-05 | 4.1493e-06 | 4.4151e-07 | 13 | 1 | 0.2 | 10 | 100.2814 | -0.46651 | -0.2545 |  | 0.082754 | 0.0088054 | 0.0010113 | 0.0019944 | 0.0019944 |  |  |  |  | -0.4313 | -0.16943 |  | -0.29593 | -0.11641
6 |  |  | 0.60465 | -1.4718e-05 | -8.2213e-06 | 1.0764e-05 | 5.2209e-07 | 1.0241e-07 | 12 | 1 | 0.2 | 10 | 60.4649 | -0.48682 | -0.27194 |  | 0.017269 | 0.0033873 | 0.0017801 | 0.0033077 | 0.0033077 |  |  |  |  | -0.46773 | -0.16443 |  | -0.019093 | -0.10751
7 |  |  | 0.4003 | -1.1078e-05 | -4.8904e-06 | 7.45e-07 | 3.363e-10 | 3.5575e-20 | 12 | 1 | 0.2 | 10 | 40.0301 | -0.5535 | -0.24434 |  | 1.6803e-05 | 1.7774e-15 | 0.00018611 | 0.0049962 | 0.0049962
8 |  |  | 0.2004 | -6.0135e-06 | -2.0355e-06 | 1.7814e-07 | 4.5007e-08 | 5.3271e-08 | 11 | 1 | 0.2 | 10 | 20.0404 | -0.60014 | -0.20314 |  | 0.0044916 | 0.0053163 | 8.8892e-05 | 0.0099798 | 0.0099798
9 |  |  | 0.10062 | -3.4014e-06 | -1.0467e-06 | 5.6938e-08 | 3.3746e-08 | 8.8462e-10 | 10 | 1 | 0.2 | 10 | 10.0624 | -0.67607 | -0.20804 |  | 0.0067073 | 0.00017583 | 5.6585e-05 | 0.019876 | 0.019876
10 |  |  | 0.050611 | -1.7806e-06 | -4.843e-07 | 2.0775e-08 | 2.6067e-08 | 3.0459e-08 | 9 | 1 | 0.2 | 10 | 5.0611 | -0.70366 | -0.19138 |  | 0.010301 | 0.012036 | 4.1048e-05 | 0.039517 | 0.039517
11 |  |  | 0.010479 | -4.0284e-07 | -1.0502e-07 | 5.5172e-08 | 4.6534e-08 | 2.9096e-08 | 8 | 1 | 0.2 | 10 | 1.0479 | -0.76884 | -0.20043 |  | 0.088811 | 0.055531 | 0.00052648 | 0.19085 | 0.19085 |  | Do a 1% point
12 |  |  | 0.010501 | -4.1343e-07 | -1.0579e-07 | 4.9333e-08 | 3.4753e-08 | 3.784e-08 | 7 | 2 | 0.2 | 10 | 1.0501 | -0.78739 | -0.20147 |  | 0.06619 | 0.072069 | 0.00046979 | 0.19046 | 0.19046 |  | Change reserve to Low Noise at 1% point
13 |  |  | 1.2515 | -4.5118e-05 | -1.7602e-05 | 3.7665e-05 | 6.9747e-06 | 1.3263e-06 | 13 | 1 | 0.2 | 10 | 125.1455 | -0.72106 | -0.2813 |  | 0.11147 | 0.021196 | 0.0030097 | 0.0015981 | 0.0015981 |  | Test 125%. 3458A on 10V range
14 |  |  | 1.2466 | -3.8538e-05 | -1.6545e-05 | 7.7463e-06 | 6.2534e-06 | 8.8737e-07 | 13 | 1 | 0.2 | 10 | 124.6603 | -0.61829 | -0.26545 |  | 0.10033 | 0.014237 | 0.0006214 | 0.0016044 | 0.0016044 |  | Repeat 125%. 3458A on 10V range
15 |  |  | 1.2502 | -2.543e-05 | -1.5054e-05 | 1.9368e-05 | 2.4174e-06 | 3.555e-07 | 13 | 1 | 0.2 | 10 | 125.0151 | -0.40683 | -0.24084 |  | 0.038674 | 0.0056872 | 0.0015492 | 0.0015998 | 0.0015998 |  | Repeat 125%. 3458A on 10V range
16 |  |  | 1.2227 | -2.7013e-05 | -1.4657e-05 | 3.4643e-05 | 4.4635e-06 | 5.2247e-07 | 13 | 1 | 0.2 | 10 | 122.2659 | -0.44187 | -0.23976 |  | 0.073013 | 0.0085465 | 0.0028334 | 0.0016358 | 0.0016358 |  | Repeat 120%. 3458A on 10V range
17
18 | 2018-06-27 | 47 Hz zero burden on aux sec | 1.1959 | -1.8629e-05 | -1.5758e-05 | 2.0512e-05 | 1.184e-06 | 2.3491e-09 | 13 | 1 | 0.2 | 10 | 119.5947 | -0.31154 | -0.26352 |  | 0.019799 | 3.9284e-05 | 0.0017151 | 0.0016723 | 0.0016723 |  | Start at 120%. 3458A on 1V range.
19 |  |  | 1.014 | -1.8226e-05 | -1.4338e-05 | 7.6125e-05 | 2.9728e-07 | 1.4768e-08 | 13 | 1 | 0.2 | 10 | 101.3976 | -0.3595 | -0.28281 |  | 0.0058636 | 0.0002913 | 0.0075076 | 0.0019724 | 0.0019724
20 |  |  | 0.58491 | -1.8332e-05 | -9.0753e-06 | 2.0526e-05 | 8.8501e-07 | 3.4878e-08 | 12 | 1 | 0.2 | 10 | 58.491 | -0.62683 | -0.31031 |  | 0.030261 | 0.0011926 | 0.0035092 | 0.0034193 | 0.0034193
21 |  |  | 0.39296 | -1.4363e-05 | -6.4746e-06 | 0.00010334 | 9.515e-07 | 2.213e-07 | 12 | 1 | 0.2 | 10 | 39.2959 | -0.73102 | -0.32953 |  | 0.048427 | 0.011263 | 0.026297 | 0.0050896 | 0.0050896
22 |  |  | 0.20859 | -1.8083e-05 | -5.6377e-06 | 5.9353e-05 | 6.512e-07 | 7.2215e-08 | 12 | 1 | 0.2 | 10 | 20.8592 | -1.7338 | -0.54054 |  | 0.062438 | 0.0069241 | 0.028454 | 0.0095881 | 0.0095881 |  | Noisy
23 |  |  | 0.099812 | -5.4424e-06 | -1.4844e-06 | 2.2156e-07 | 1.9534e-08 | 3.1322e-09 | 11 | 1 | 0.2 | 10 | 9.9812 | -1.0905 | -0.29743 |  | 0.0039141 | 0.00062762 | 0.00022198 | 0.020038 | 0.020038
24 |  |  | 0.050008 | -2.8176e-06 | -7.2146e-07 | 5.5181e-08 | 1.5357e-08 | 3.5494e-08 | 10 | 1 | 0.2 | 10 | 5.0008 | -1.1269 | -0.28854 |  | 0.0061419 | 0.014195 | 0.00011034 | 0.039994 | 0.039994
25 |  |  | 0.010217 | -6.1858e-07 | -1.6485e-07 | 7.0196e-08 | 5.2625e-08 | 5.542e-08 | 9 | 1 | 0.2 | 10 | 1.0217 | -1.2109 | -0.32272 |  | 0.10302 | 0.10849 | 0.00068707 | 0.19576 | 0.19576 |  | Do a 1% point
26 |  |  | 0.010249 | -5.9999e-07 | -1.7713e-07 | 7.262e-08 | 4.8394e-08 | 4.5575e-08 | 8 | 2 | 0.2 | 10 | 1.0249 | -1.1708 | -0.34566 |  | 0.094437 | 0.088936 | 0.00070856 | 0.19514 | 0.19514 |  | Change reserve to Low Noise at 1% point
27 |  |  | 1.2425 | -1.705e-05 | -1.5497e-05 | 2.1176e-05 | 1.4416e-07 | 4.3262e-08 | 13 | 1 | 0.2 | 10 | 124.252 | -0.27445 | -0.24944 |  | 0.0023204 | 0.00069636 | 0.0017042 | 0.0016096 | 0.0016096 |  | Test 125%. 3458A on 10V range
28 |  |  | 1.2505 | -1.6865e-05 | -1.563e-05 | 2.0021e-05 | 1.6225e-09 | 5.7998e-08 | 13 | 1 | 0.2 | 10 | 125.0549 | -0.26972 | -0.24997 |  | 2.5948e-05 | 0.00092757 | 0.001601 | 0.0015993 | 0.0015993 |  | Repeat 125%. 3458A on 10V range
29 |  |  | 1.2009 | -1.6849e-05 | -1.4843e-05 | 2.0441e-05 | 4.2866e-08 | 1.9506e-07 | 13 | 1 | 0.2 | 10 | 120.0884 | -0.28061 | -0.24719 |  | 0.00071391 | 0.0032486 | 0.0017022 | 0.0016654 | 0.0016654 |  | Repeat 120%. 3458A on 10V range
30
31 |  |  | 0.20394 | -6.0525e-06 | -2.1678e-06 | 3.6579e-08 | 8.4567e-09 | 7.1096e-08 | 11 | 1 | 0.2 | 10 | 20.3939 | -0.59356 | -0.2126 |  | 0.00082934 | 0.0069723 | 1.7936e-05 | 0.0098068 | 0.0098068 |  | Repeat 20% point about 1215. everything much quieter.
32
33 | 2018-06-27 | Inner primary voltage drops (Ta)
34 |  |  |  | main sec |  |  | sec internal | total | pri only |  | V drop | cond S | relative | [9 empty] | 2018-06-27
35 |  |  | shunt V | burden V | volt drop | shunt | burden V | sec V | V drop | section | ea section | ea section | current | [9 empty] | Shunt V | Burden V | Section V | Shunt V
36 |  | Volt drop section 1 | 0.99944 | 1.183 | 1.1186 | 0.2 | 1.3392 | 2.5223 | 0.48807 | 1 | 0.48807 | 10.2387 | 1.0255 | [9 empty] | 0.9994 | 1.183 | 1.1186 | 0.99947 |  |  | Burden V seems high with respect to history.
37 |  | Volt drop sections 1+2 | 0.99946 | 1.1838 | 2.2473 | 0.2 | 1.3393 | 2.5231 | 0.98573 | 2 | 0.49766 | 10.0416 | 1.0058 | [9 empty] | 0.99945 | 1.1838 | 1.1287 | 0.99947
38 |  | Volt drop sections 1+2+3 | 0.99947 | 1.1835 | 3.3891 | 0.2 | 1.3393 | 2.5228 | 1.497 | 3 | 0.51122 | 9.7753 | 0.9791 | [9 empty] | 0.99947 | 1.1835 | 1.1418 | 0.99947 |  |  | The shunt was connected through a terminal strip, connect directly to the main secondary terminals.
39 |  | Volt drop sections 1+2+3+4 | 0.99947 | 1.1835 | 4.5256 | 0.2 | 1.3393 | 2.5228 | 2.0027 | 4 | 0.50579 | 9.8803 | 0.98961 | [9 empty] | 0.99947 | 1.1835 | 1.1365 | 0.99947
40 |  |  |  |  | These are the sections added together. See column 'S'. Hand waving exercise.
41
42 | 2018-07-11 | Inner primary voltage drops (Ta) repeated
43 |  |  |  | main sec |  |  | sec internal | total | pri only |  | V drop | cond S | relative | [15 empty] | Clean series network connections.
44 |  |  | shunt V | burden V | volt drop | shunt | burden V | sec V | V drop | section | ea section | ea section | current | [9 empty] | Shunt V | Burden V | Section V | Shunt V
45 |  | Volt drop section 1 | 1.0005 | 1.1105 | 1.0595 | 0.2 | 1.3407 | 2.4512 | 0.44673 | 1 | 0.44673 | 11.1983 | 1.0281 | [9 empty] | 1.0008 | 1.1105 | 1.0595 | 1.0003
46 |  | Volt drop sections 1+2 | 1.0006 | 1.1085 | 2.1354 | 0.2 | 1.3408 | 2.4493 | 0.91077 | 2 | 0.46404 | 10.781 | 0.98976 | [9 empty] | 1.0008 | 1.1085 | 1.0759 | 1.0003
47 |  | Volt drop sections 1+2+3 | 1.0001 | 1.1109 | 3.2171 | 0.2 | 1.3401 | 2.4509 | 1.3789 | 3 | 0.46817 | 10.6804 | 0.98053 | [9 empty] | 1.0001 | 1.1109 | 1.0817 | 1
48 |  | Volt drop sections 1+2+3+4 | 1.0007 | 1.1119 | 4.2905 | 0.2 | 1.341 | 2.4529 | 1.8376 | 4 | 0.45862 | 10.9105 | 1.0016 | [9 empty] | 1.0009 | 1.1119 | 1.0734 | 1.0006
49 |  |  |  |  | These are the sections added together. See column 'S'. Hand waving exercise.
50
51 | 2018-07-11 | Voltage drops, middle primary of Ta |  |  |  |  | sec internal | total | pri only |  | V drop | cond S | relative | [13 empty] | Cumulative
52 |  |  |  |  | sec re 1 V | shunt | burden V | sec V | V drop | section | ea section | ea section | current | [9 empty] | Shunt V | Section V | Shunt V |  | section V
53 | Repeated | Volt drop section 1 | 1.0005 | 0.35453 | 0.35436 | 0.2 | 1.34 | 2.34 | 0.23736 | 1 | 0.23736 | 4.213 | 1.0644 | [9 empty] | 1.0007 | 0.35453 | 1.0003 |  | 0.35453
54 | below. | Volt drop sections 1+2 | 1.0008 | 0.7415 | 0.74094 | 0.2 | 1.34 | 2.34 | 0.50694 | 2 | 0.26958 | 3.7095 | 0.93721 | [9 empty] | 1.0009 | 0.38697 | 1.0006 |  | 0.7415
55 | on 25th | Volt drop sections 1+2+3 | 1.0003 | 1.1303 | 1.13 | 0.2 | 1.34 | 2.34 | 0.77895 | 3 | 0.27202 | 3.6762 | 0.92881 | [9 empty] | 1.0005 | 0.38882 | 1.0002 |  | 1.1303
56 |  | Volt drop sections 1+2+3+4 | 1.0006 | 1.484 | 1.4832 | 0.2 | 1.34 | 2.34 | 1.0152 | 4 | 0.23622 | 4.2333 | 1.0696 | [9 empty] | 1.0009 | 0.35372 | 1.0003 |  | 1.484
57
58
59 | 2018-07-11 | Voltage drops,outer primary of Ta |  | measure across each section in turn, do not include any link | [6 empty] | section resistance | section g | Ik/Ips | [9 empty] | Shunt V | Section V | Shunt V
60 |  | section 1 | 1.0001 | 0.16112 | 0.16111 | 0.2 | 1.34 | 2.34 | 0.13771 | 1 | 0.0013769 | 726.2506 | 1.1482 | [9 empty] | 1.0009 | 0.16112 | 0.99931
61 |  | section 2 | 1.0004 | 0.18025 | 0.18018 | 0.2 | 1.34 | 2.34 | 0.15678 | 2 | 0.0015673 | 638.0589 | 1.0088 | [9 empty] | 1.0009 | 0.18025 | 0.99982
62 |  | section 3 | 1.0005 | 0.17938 | 0.17929 | 0.2 | 1.34 | 2.34 | 0.15589 | 3 | 0.0015582 | 641.762 | 1.0146 | [9 empty] | 1.0008 | 0.17938 | 1.0001
63 |  | section 4 | 1.0005 | 0.22671 | 0.22659 | 0.2 | 1.34 | 2.34 | 0.20319 | 4 | 0.0020309 | 492.4012 | 0.7785 | [9 empty] | 1.001 | 0.22671 | 1
64 |  | section 5 | 1.0004 | 0.17413 | 0.17406 | 0.2 | 1.34 | 2.34 | 0.15066 | 5 | 0.0015059 | 664.0432 | 1.0499 | [9 empty] | 1.0006 | 0.17413 | 1.0002
65
66 |  | Retest using 50A, not so much heating. Section V, test to the washer just behind the series network.
67 | 2018-07-12 | Voltage drops,outer primary of Ta |  | measure across each section in turn, do not include any link | [6 empty] | section resistance | section g | Ik/Ips | [9 empty] | Shunt V | Section mV | Shunt V
68 |  | section 1 | 0.50106 | 0.072064 | 0.14382 | 0.2 | 1.34 | 2.34 | 0.12042 | 1 | 0.0024034 | 416.0837 | 1.1273 | [9 empty] | 0.50112 | 72.0639 | 0.501
69 |  | section 2 | 0.50084 | 0.077295 | 0.15433 | 0.2 | 1.34 | 2.34 | 0.13093 | 2 | 0.0026142 | 382.5228 | 1.0363 | [9 empty] | 0.5009 | 77.2953 | 0.50079
70 |  | section 3 | 0.50059 | 0.079708 | 0.15923 | 0.2 | 1.34 | 2.34 | 0.13583 | 3 | 0.0027134 | 368.5376 | 0.99844 | [9 empty] | 0.50067 | 79.7082 | 0.5005
71 |  | section 4 | 0.50025 | 0.10101 | 0.20191 | 0.2 | 1.34 | 2.34 | 0.17851 | 4 | 0.0035685 | 280.2316 | 0.7592 | [9 empty] | 0.5003 | 101.0072 | 0.5002
72 |  | section 5 | 0.50012 | 0.074516 | 0.149 | 0.2 | 1.34 | 2.34 | 0.1256 | 5 | 0.0025114 | 398.1918 | 1.0788 | [9 empty] | 0.50016 | 74.5158 | 0.50007
73
74 | 2018-07-12 | Volt Drops, 20/100 A pri of Tb |  |  |  |  | sec internal | total | pri only |  | V drop | cond S | relative | [13 empty] | Cumulative
75 |  |  | shunt V | sec drop | sec re 1 V | shunt | burden V | sec V | V drop | section | ea section | ea section | current | [9 empty] | Shunt V | Section V | Shunt V |  | section V
76 |  | Volt drop section 1 | 0.50007 | 0.21032 | 0.42059 | 0.2 | 1.585 | 2.585 | 0.29134 | 1 | 0.29134 | 3.4325 | 0.90999 | [9 empty] | 0.50006 | 0.21032 | 0.50008 |  | 0.21032
77 |  | Volt drop sections 1+2 | 0.50018 | 0.4151 | 0.8299 | 0.2 | 1.585 | 2.585 | 0.5714 | 2 | 0.28007 | 3.5706 | 0.94661 | [9 empty] | 0.50014 | 0.20478 | 0.50021 |  | 0.4151
78 |  | Volt drop sections 1+2+3 | 0.50032 | 0.62125 | 1.2417 | 0.2 | 1.585 | 2.585 | 0.85395 | 3 | 0.28255 | 3.5392 | 0.93828 | [9 empty] | 0.50034 | 0.20615 | 0.50031 |  | 0.62125
79 |  | Volt drop sections 1+2+3+4 | 0.50031 | 0.80537 | 1.6097 | 0.2 | 1.585 | 2.585 | 1.0927 | 4 | 0.23877 | 4.1881 | 1.1103 | [9 empty] | 0.50033 | 0.18412 | 0.5003 |  | 0.80537
80 |  | Volt drop sections 1+2+3+4+5 | 0.50078 | 0.99211 | 1.9811 | 0.2 | 1.585 | 2.585 | 1.3349 | 5 | 0.24215 | 4.1296 | 1.0948 | [9 empty] | 0.50083 | 0.18674 | 0.50073 |  | 0.99211
81
82 | [6 empty] | sec internal | total | pri only |  | V drop | cond S | relative | [13 empty] | Cumulative
83 | 2018-07-12 | Volt drops, outer primary of Tb |  |  | sec re 1 V | shunt | burden V | sec V | V drop | section | ea section | ea section | current | [9 empty] | Shunt V | Section mV | Shunt V |  | section V
84 |  | Volt drop section 1 | 0.50003 | 0.09926 | 0.19851 | 0.2 | 1.585 | 2.585 | 0.069257 | 1 | 0.069257 | 14.439 | 1.0406 | [9 empty] | 0.50014 | 99.2599 | 0.49992 |  | 99.2599
85 |  | Volt drop sections 1+2 | 0.50152 | 0.20205 | 0.40288 | 0.2 | 1.585 | 2.585 | 0.14438 | 2 | 0.075125 | 13.3111 | 0.95936 | [9 empty] | 0.50169 | 102.7944 | 0.50135 |  | 202.0543
86
87
88 |  | COUPLING EQUALITY OF INNER PRIMARY of Ta |  |  |  | main secondary open circuit. aux sec to common resistor. test current 5 A monitored with Tb, 20/5 | [11 empty] | Use twisted pair from SR830 to resistor. Coaxial cables couple together unpredictably.
89 |  | pri section 1 vs 2: effect of earthing |  |  |  | Tb Aux sec shorted. Section 1 is at Low end of Ta. (Low Potential) which is NOT the "L" end, that "L" stands for "live" on the CT. (BM) | [16 empty] | Test Point Data. Copy/past into spreadsheet.
90 |  |  | Shunt | X | Y | Sh V | X stdev | Y stdev | CT ratio | Shunt | Com R | test current | Aux sec I |  | N(m2-m1)/Z1 |  |  | dm x dI | dm x dI | S/P error |  |  | SD ppm | SD ppm | SD ppm | noise sd | noise sd |  |  |  |  | IVD
91 |  |  | Volts | V | V | stdev/100 | V | V |  |  |  |  | Re uA | Im uA | Re ppm | Im ppm | 1-Ipk/Ips |  |  | Re | Im |  | X | Y | shunt V | X | Y |  |  |  |  | setting
92 | 2018-07-16 | pri floating | 0.24983 | -9.9057e-07 | -4.3354e-06 | 1.2034e-06 | 7.2889e-08 | 3.2347e-08 | 4 | 0.2 | 10 | 4.9966 | -0.099057 | -0.43354 | -0.079299 | -0.34706 | [6 empty] | 0.24983 | -9.9057e-07 | -4.3354e-06 | 1.2034e-06 | 7.2889e-08 | 3.2347e-08 | 10 | 1
93 |  | pri LP earthed | 0.24961 | -8.6932e-07 | -4.6132e-06 | 1.4876e-06 | 3.5961e-08 | 4.7491e-08 | 4 | 0.2 | 10 | 4.9922 | -0.086932 | -0.46132 | -0.069654 | -0.36963 | [6 empty] | 0.24961 | -8.6932e-07 | -4.6132e-06 | 1.4876e-06 | 3.5961e-08 | 4.7491e-08 | 10 | 1
94 |  | pri centre earthed | 0.24919 | -8.9588e-07 | -4.7433e-06 | 3.9719e-07 | 1.9007e-08 | 2.3665e-08 | 4 | 0.2 | 10 | 4.9838 | -0.089588 | -0.47433 | -0.071903 | -0.3807 | [6 empty] | 0.24919 | -8.9588e-07 | -4.7433e-06 | 3.9719e-07 | 1.9007e-08 | 2.3665e-08 | 10 | 1
95 |  | pri HP earthed | 0.24966 | -1.0245e-06 | -4.2419e-06 | 8.0938e-07 | 4.1658e-08 | 3.1539e-08 | 4 | 0.2 | 10 | 4.9932 | -0.10245 | -0.42419 | -0.082069 | -0.33982 | [6 empty] | 0.24966 | -1.0245e-06 | -4.2419e-06 | 8.0938e-07 | 4.1658e-08 | 3.1539e-08 | 10 | 1
96
97 |  | pri centre Wagnered to earth | 0.24927 | -9.4924e-07 | -4.3728e-06 | 2.0109e-06 | 4.2188e-08 | 2.7717e-08 | 4 | 0.2 | 10 | 4.9854 | -0.094924 | -0.43728 | -0.076163 | -0.35085 | [6 empty] | 0.24927 | -9.4924e-07 | -4.3728e-06 | 2.0109e-06 | 4.2188e-08 | 2.7717e-08 | 10 | 1
98 |  | Wagnered, 100k across section 1 | 0.24947 | 1.0523e-05 | -7.0106e-08 | 8.7824e-07 | 2.2384e-08 | 2.8354e-08 | 4 | 0.2 | 10 | 4.9893 | 1.0523 | -0.0070106 | 0.84367 | -0.0056205 | [6 empty] | 0.24947 | 1.0523e-05 | -7.0106e-08 | 8.7824e-07 | 2.2384e-08 | 2.8354e-08 | 11 | 1
99 |  | Intercompare sections with pri centre Wagnered to earth
100 |  | sec 1 vs sec 2 | 0.24907 | -9.3459e-07 | -4.3555e-06 | 1.8772e-06 | 1.9778e-08 | 1.531e-08 | 4 | 0.2 | 10 | 4.9815 | -0.093459 | -0.43555 | -0.075045 | -0.34973 | 0.01024 | -0.00076845 | -0.0035812 |  |  |  | 0.24907 | -9.3459e-07 | -4.3555e-06 | 1.8772e-06 | 1.9778e-08 | 1.531e-08 | 10 | 1 |  | 665
101 |  | sec 1 vs sec 3 | 0.25044 | 1.167e-07 | 1.563e-06 | 1.7656e-06 | 2.64e-08 | 2.1381e-08 | 4 | 0.2 | 10 | 5.0088 | 0.01167 | 0.1563 | 0.0093196 | 0.12482 | 0.019472 | 0.00018147 | 0.0024306 |  |  |  | 0.25044 | 1.167e-07 | 1.563e-06 | 1.7656e-06 | 2.64e-08 | 2.1381e-08 | 9 | 1 |  | 668
102 |  | sec 1 vs sec 4 | 0.24943 | -1.5568e-06 | -2.6296e-06 | 6.0643e-07 | 2.888e-08 | 3.4954e-08 | 4 | 0.2 | 10 | 4.9886 | -0.15568 | -0.26296 | -0.12483 | -0.21085 | -0.0016455 | 0.0002054 | 0.00034694 | -0.00038158 | -0.00080369 |  | 0.24943 | -1.5568e-06 | -2.6296e-06 | 6.0643e-07 | 2.888e-08 | 3.4954e-08 | 10 | 1 |  | 664
103 |  | sec 2 vs sec 3 | 0.25024 | 1.2317e-06 | 5.6897e-06 | 1.7535e-06 | 3.5465e-08 | 3.876e-08 | 4 | 0.2 | 10 | 5.0047 | 0.12317 | 0.56897 | 0.098439 | 0.45475 |  |  |  | -0.0028875 | -0.00067812 |  | 0.25024 | 1.2317e-06 | 5.6897e-06 | 1.7535e-06 | 3.5465e-08 | 3.876e-08 | 11 | 1 |  | 660
104 |  | sec 2 vs sec 4 | 0.25024 | -5.7064e-07 | 1.4461e-06 | 1.9985e-06 | 4.5336e-08 | 4.6981e-08 | 4 | 0.2 | 10 | 5.0047 | -0.057064 | 0.14461 | -0.045608 | 0.11558 | [6 empty] | 0.25024 | -5.7064e-07 | 1.4461e-06 | 1.9985e-06 | 4.5336e-08 | 4.6981e-08 | 9 | 1 |  | 672
105 |  | sec 3 vs sec 4 | 0.25006 | -1.5534e-06 | -4.3086e-06 | 2.6223e-06 | 1.7692e-08 | 1.8708e-08 | 4 | 0.2 | 10 | 5.0011 | -0.15534 | -0.43086 | -0.12424 | -0.34461 | [6 empty] | 0.25006 | -1.5534e-06 | -4.3086e-06 | 2.6223e-06 | 1.7692e-08 | 1.8708e-08 | 10 | 1 |  | 672
106
107 |  | Magnetic coupling equality, 20/100 A pri of Tb | [6 empty] | CT ratio | Shunt | Com R | test current | Aux sec I |  | N(m2-m1)/Z1 |  |  | dm x dI | dm x dI | S/P error
108 | [12 empty] | Re uA | Im uA | Re ppm | Im ppm | 1-Ipk/Ips | Re | Im | Re | Im
109 |  | section 1 vs 2:wagner | 1.0115 | -1.2434e-06 | -5.2321e-06 | 9.3489e-06 | 2.3099e-08 | 2.9481e-08 | 4 | 0.2 | 10 | 20.2297 | -0.12434 | -0.52321 | -0.024586 | -0.10345 | [6 empty] | 1.0115 | -1.2434e-06 | -5.2321e-06 | 9.3489e-06 | 2.3099e-08 | 2.9481e-08 | 10 | 1
110 |  | sec 1 vs 2: 100k across 1 | 1.0016 | 2.3726e-07 | -4.9572e-06 | 5.6101e-06 | 2.817e-08 | 2.5207e-08 | 4 | 0.2 | 10 | 20.0324 | 0.023726 | -0.49572 | 0.0047375 | -0.098983 | [6 empty] | 1.0016 | 2.3726e-07 | -4.9572e-06 | 5.6101e-06 | 2.817e-08 | 2.5207e-08 | 10 | 1
111 |  | sec 1 vs sec 2: wagner offset | 0.99723 | -1.2267e-06 | -5.196e-06 | 3.3258e-06 | 2.1269e-08 | 4.1069e-08 | 4 | 0.2 | 10 | 19.9446 | -0.12267 | -0.5196 | -0.024603 | -0.10421 | [6 empty] | 0.99723 | -1.2267e-06 | -5.196e-06 | 3.3258e-06 | 2.1269e-08 | 4.1069e-08 | 10 | 1
112 |  | Offset Wagner from 8230000 to8430000
113 |  | sec 1 vs sec 2 | 0.99483 | -1.2272e-06 | -5.2001e-06 | 1.3454e-06 | 1.5934e-08 | 2.399e-08 | 4 | 0.2 | 10 | 19.8965 | -0.12272 | -0.52001 | -0.024671 | -0.10454 | 0.053392 | -0.0013172 | -0.0055817 |  |  |  | 0.99483 | -1.2272e-06 | -5.2001e-06 | 1.3454e-06 | 1.5934e-08 | 2.399e-08 | 10 | 1
114 |  | sec 1 vs sec3 | 0.9945 | -1.1631e-06 | -7.8217e-06 | 1.9467e-05 | 3.818e-08 | 4.5535e-08 | 4 | 0.2 | 10 | 19.89 | -0.11631 | -0.78217 | -0.023391 | -0.1573 | 0.06172 | -0.0014437 | -0.0097085 |  |  |  | 0.9945 | -1.1631e-06 | -7.8217e-06 | 1.9467e-05 | 3.818e-08 | 4.5535e-08 | 11 | 1
115 |  | sec 1 vs sec 4 | 1.0004 | -3.9949e-07 | -7.2568e-06 | 1.7008e-05 | 4.865e-08 | 4.0374e-08 | 4 | 0.2 | 10 | 20.008 | -0.039949 | -0.72568 | -0.0079866 | -0.14508 | -0.11031 | 0.00088098 | 0.016003 |  |  |  | 1.0004 | -3.9949e-07 | -7.2568e-06 | 1.7008e-05 | 4.865e-08 | 4.0374e-08 | 11 | 1
116 |  | sec 1 vs sec 5 | 1.0022 | -3.3578e-07 | -3.932e-08 | 1.5447e-05 | 2.3608e-08 | 2.8024e-08 | 4 | 0.2 | 10 | 20.0446 | -0.033578 | -0.003932 | -0.0067007 | -0.00078465 | -0.094811 | 0.0006353 | 7.4394e-05 | -0.0012446 | 0.00078735 |  | 1.0022 | -3.3578e-07 | -3.932e-08 | 1.5447e-05 | 2.3608e-08 | 2.8024e-08 | 7 | 1
117 |  | repeat test, 1 vs 5 | 0.98583 | -3.222e-07 | -1.1103e-08 | 1.1869e-05 | 2.3483e-08 | 2.9471e-08 | 4 | 0.2 | 10 | 19.7166 | -0.03222 | -0.0011103 | -0.0065367 | -0.00022524 |  |  |  | -0.0033987 | -0.00072971 |  | 0.98583 | -3.222e-07 | -1.1103e-08 | 1.1869e-05 | 2.3483e-08 | 2.9471e-08 | 7 | 1
118
119 |  | Magnetic coupling equality, middle primary of Ta | [6 empty] | CT ratio | Shunt | Com R | test current | Aux sec I |  | N(m2-m1)/Z1 |  |  | dm x dI | dm x dI | S/P error
120 |  | Read this with Ta and Tb LP terminals in series. Phase on CH set to 0° so that the SR830 phase is correcting about 0°. | [10 empty] | Re uA | Im uA | Re ppm | Im ppm | 1-Ipk/Ips | Re | Im | Re | Im | [10 empty] | repeat
121 | Repeated | sec 1 vs sec 2 | 1.0094 | 9.1214e-07 | 9.3169e-07 | 1.0998e-05 | 3.7525e-08 | 3.0629e-08 | 4 | 0.2 | 10 | 20.1883 | 0.091214 | 0.093169 | 0.018073 | 0.01846 | [6 empty] | 1.0094 | 9.1214e-07 | 9.3169e-07 | 1.0998e-05 | 3.7525e-08 | 3.0629e-08 | 9 | 1 |  | 1.0307 | 8.9681e-07 | 2.6312e-06 | 1.1654e-05 | 2.2275e-08 | 1.5655e-08 | 10 | 1
122 | below. | sec 1 vs 2: 100k across 1 | 0.99598 | 1.8603e-06 | 1.2312e-06 | 5.0417e-06 | 3.8075e-08 | 3.4627e-08 | 4 | 0.2 | 10 | 19.9196 | 0.18603 | 0.12312 | 0.037356 | 0.024723 | [6 empty] | 0.99598 | 1.8603e-06 | 1.2312e-06 | 5.0417e-06 | 3.8075e-08 | 3.4627e-08 | 9 | 1 |  | 1.0126 | 1.9582e-06 | 3.014e-06 | 4.5192e-06 | 3.0404e-08 | 3.3112e-08 | 10 | 1
123 | on 25th
124 |  | sec 1 vs sec 2 | 1.0011 | 9.2361e-07 | 9.971e-07 | 4.6644e-06 | 3.0295e-08 | 2.6999e-08 | 4 | 0.2 | 10 | 20.0228 | 0.092361 | 0.09971 | 0.018451 | 0.019919 | 0.062789 | 0.0011585 | 0.0012507 |  |  |  | 1.0011 | 9.2361e-07 | 9.971e-07 | 4.6644e-06 | 3.0295e-08 | 2.6999e-08 | 9 | 1 |  | 1.0059 | 9.1582e-07 | 1.7973e-06 | 2.5351e-06 | 3.5977e-08 | 2.7769e-08 | 9 | 1
125 |  | sec 1 vs sec3 | 1.007 | 1.807e-06 | 3.0976e-06 | 1.6245e-05 | 5.0866e-08 | 6.1196e-08 | 4 | 0.2 | 10 | 20.1399 | 0.1807 | 0.30976 | 0.035888 | 0.061521 | 0.071194 | 0.002555 | 0.0043799 |  |  |  | 1.007 | 1.807e-06 | 3.0976e-06 | 1.6245e-05 | 5.0866e-08 | 6.1196e-08 | 10 | 1 |  | 1.0049 | 1.5973e-06 | 3.1889e-06 | 1.6597e-05 | 2.7058e-08 | 5.895e-08 | 10 | 1
126 |  | sec 1 vs sec 4 | 1.0101 | -7.4207e-07 | -2.0721e-06 | 5.1378e-06 | 4.9237e-08 | 6.5255e-08 | 4 | 0.2 | 10 | 20.2014 | -0.074207 | -0.20721 | -0.014693 | -0.041029 | -0.069554 | 0.001022 | 0.0028537 | 0.0047355 | 0.0084844 |  | 1.0101 | -7.4207e-07 | -2.0721e-06 | 5.1378e-06 | 4.9237e-08 | 6.5255e-08 | 10 | 1 |  | 1.0265 | -6.5037e-07 | -3.1807e-06 | 0.00013231 | 7.144e-08 | 6.5367e-08 | 10 | 1
127 | [19 empty] | 0.001573 | 0.00069226
128 | [19 empty] | 0.0048877 | -0.0034944
129 |  | Magnetic coupling equality, outer primary of Ta | [6 empty] | CT ratio | Shunt | Com R | test current | Aux sec I |  | N(m2-m1)/Z1 |  |  | dm x dI | dm x dI | S/P error
130 | [12 empty] | Re uA | Im uA | Re ppm | Im ppm | Ipk/Ips | Re | Im | Re | Im
131 | 2018-07-17 | sec 1 vs sec 2 | 0.99586 | -3.4797e-06 | -7.9654e-06 | 4.8641e-06 | 2.5315e-08 | 7.4818e-08 | 20 | 0.2 | 10 | 99.5862 | -0.34797 | -0.79654 | -0.069884 | -0.15997 | [6 empty] | 0.99586 | -3.4797e-06 | -7.9654e-06 | 4.8641e-06 | 2.5315e-08 | 7.4818e-08 | 11 | 1
132 |  | 27k across sec 1 | 0.99155 | -2.9568e-06 | -7.9598e-06 | 4.0572e-06 | 4.1685e-09 | 5.6665e-08 | 20 | 0.2 | 10 | 99.1546 | -0.29568 | -0.79598 | -0.05964 | -0.16055 | [6 empty] | 0.99155 | -2.9568e-06 | -7.9598e-06 | 4.0572e-06 | 4.1685e-09 | 5.6665e-08 | 11 | 1
133 | These are | 27k across sec 2 | 0.98811 | -4.0098e-06 | -8.1237e-06 | 3.5359e-06 | 6.1934e-08 | 1.3794e-07 | 20 | 0.2 | 10 | 98.8112 | -0.40098 | -0.81237 | -0.08116 | -0.16443 | [6 empty] | 0.98811 | -4.0098e-06 | -8.1237e-06 | 3.5359e-06 | 6.1934e-08 | 1.3794e-07 | 11 | 1
134 | repeated
135 | below. | [10 empty] | Applied | Aux sec current, uA |  | N(m2-m1)/Z1 |  |  | dm x dI | dm x dI | S/P correction |  |  | 1.0081 | -9.8166e-07 | -2.2967e-06 | 9.9047e-06 | 1.3302e-08 | 3.6145e-08 | 10 | 1 | 1-2 | -3.9 | °
136 |  | Magnetic coupling equality, outer primary of Ta | [6 empty] | CT ratio | Shunt | Aux Sec res | current | Re | Im | Re ppm | Im ppm | Ipk/Ips |  |  | Re | Im |  | 1.0074 | -4.8254e-06 | -1.654e-05 | 4.5944e-06 | 3.7523e-08 | 8.1677e-08 | 12 | 1 | 2-1 | -3.9 | ° | reverse
137 |  | sec 1 vs sec 2 | 1.0063 | -1.5642e-06 | -2.4419e-06 | 5.0067e-06 | 3.4583e-08 | 1.1806e-08 | 20 | 0.2 | 10 | 100.5085 | -0.35925 | -0.88147 | -0.071486 | -0.1754 | -0.036328 | 0.002597 | 0.0063721 |  |  |  | 1.0063 | -1.5642e-06 | -2.4419e-06 | 5.0067e-06 | 3.4583e-08 | 1.1806e-08 | 10 | 1 | 1-2 | 0 | °
138 |  | sec 1 vs sec 2 reverse | 1.0038 | -5.6207e-06 | -1.5188e-05 | 3.5708e-06 | 2.5624e-08 | 4.3517e-08 | [14 empty] | 1.0038 | -5.6207e-06 | -1.5188e-05 | 3.5708e-06 | 2.5624e-08 | 4.3517e-08 | 12 | 1 | 2-1 | 0 | ° | reverse
139 |  | sec 1 vs sec3 | 0.99921 | 2.2491e-06 | 7.837e-06 | 5.6206e-06 | 4.5243e-08 | 5.3659e-09 | 20 | 0.2 | 10 | 100.3315 | -0.43245 | -0.92561 | -0.086204 | -0.18451 | 0.0015602 | -0.0001345 | -0.00028787 |  |  |  | 0.99921 | 2.2491e-06 | 7.837e-06 | 5.6206e-06 | 4.5243e-08 | 5.3659e-09 | 11 | 1
140 |  | sec 1 vs sec3 reverse | 1.0074 | -1.0898e-05 | -2.6349e-05 | 5.8827e-06 | 3.508e-09 | 4.141e-10 | [14 empty] | 1.0074 | -1.0898e-05 | -2.6349e-05 | 5.8827e-06 | 3.508e-09 | 4.141e-10 | 13 | 1
141 |  | sec 1 vs sec 4 | 1.0034 | 6.2603e-06 | 1.5267e-05 | 7.579e-06 | 7.4415e-09 | 3.3881e-21 | 20 | 0.2 | 10 | 100.4004 | -0.45877 | -1.1859 | -0.091387 | -0.23624 | 0.2408 | -0.022006 | -0.056885 |  |  |  | 1.0034 | 6.2603e-06 | 1.5267e-05 | 7.579e-06 | 7.4415e-09 | 3.3881e-21 | 12 | 1
142 |  | sec 1 vs sec 4 reverse | 1.0046 | -1.5436e-05 | -3.8985e-05 | 6.2676e-06 | 3.9751e-08 | 2.1217e-07 | [14 empty] | 1.0046 | -1.5436e-05 | -3.8985e-05 | 6.2676e-06 | 3.9751e-08 | 2.1217e-07 | 13 | 1
143 |  | sec 1 vs sec 5 | 1.0014 | 5.5507e-06 | 1.6663e-05 | 6.5596e-06 | 2.0059e-08 | 3.2136e-10 | 20 | 0.2 | 10 | 100.2799 | -0.40787 | -0.97367 | -0.081346 | -0.19419 | -0.078779 | 0.0064083 | 0.015298 | -0.013135 | -0.035503 |  | 1.0014 | 5.5507e-06 | 1.6663e-05 | 6.5596e-06 | 2.0059e-08 | 3.2136e-10 | 12 | 1
144 |  | sec 1 vs sec 5 reverse | 1.0042 | -1.3708e-05 | -3.6136e-05 | 5.4549e-06 | 1.7923e-07 | 7.6e-10 | [11 empty] | -0.00030375 | -0.0015001 |  | 1.0042 | -1.3708e-05 | -3.6136e-05 | 5.4549e-06 | 1.7923e-07 | 7.6e-10 | 13 | 1
145 |  | sec 2 vs sec 3 | 1.0005 | 2.4635e-06 | 5.8963e-06 | 5.8887e-06 | 3.0266e-09 | 1.3004e-08 | 20 | 0.2 | 10 | 100.1998 | -0.36752 | -0.96753 | -0.073357 | -0.19312 | [6 empty] | 1.0005 | 2.4635e-06 | 5.8963e-06 | 5.8887e-06 | 3.0266e-09 | 1.3004e-08 | 11 | 1
146 |  | sec 2 vs sec 3 reverse | 1.0035 | -9.8139e-06 | -2.5247e-05 | 6.6265e-06 | 1.9095e-08 | 8.2272e-10 | [14 empty] | 1.0035 | -9.8139e-06 | -2.5247e-05 | 6.6265e-06 | 1.9095e-08 | 8.2272e-10 | 13 | 1
147 |  | sec 2 vs sec 4 | 1.0016 | 8.6669e-06 | 1.8356e-05 | 6.0436e-06 | 7.5288e-08 | 5.2084e-08 | 20 | 0.2 | 10 | 100.4418 | -0.30856 | -0.88113 | -0.061441 | -0.17545 | [6 empty] | 1.0016 | 8.6669e-06 | 1.8356e-05 | 6.0436e-06 | 7.5288e-08 | 5.2084e-08 | 12 | 1
148 |  | sec 2 vs sec 4 reverse | 1.0072 | -1.4838e-05 | -3.5979e-05 | 6.6242e-06 | 1.0461e-07 | 9.9598e-08 | [14 empty] | 1.0072 | -1.4838e-05 | -3.5979e-05 | 6.6242e-06 | 1.0461e-07 | 9.9598e-08 | 13 | 1
149 |  | sec 2 vs sec 5 | 1.0037 | 6.3807e-06 | 1.7972e-05 | 6.5176e-06 | 3.917e-08 | 1.8438e-08 | 20 | 0.2 | 10 | 100.3206 | -0.33617 | -0.89875 | -0.06702 | -0.17918 | [6 empty] | 1.0037 | 6.3807e-06 | 1.7972e-05 | 6.5176e-06 | 3.917e-08 | 1.8438e-08 | 12 | 1
150 |  | sec 2 vs sec 5 reverse | 1.0027 | -1.3104e-05 | -3.5947e-05 | 5.7948e-06 | 7.2e-10 | 8.8687e-08 | [14 empty] | 1.0027 | -1.3104e-05 | -3.5947e-05 | 5.7948e-06 | 7.2e-10 | 8.8687e-08 | 13 | 1
151 |  | sec 3 vs sec 4 | 1.0059 | 3.4588e-06 | 6.8525e-06 | 6.5521e-06 | 1.958e-08 | 4.0156e-09 | 20 | 0.2 | 10 | 100.496 | -0.37169 | -1.03 | -0.073972 | -0.20498 | [6 empty] | 1.0059 | 3.4588e-06 | 6.8525e-06 | 6.5521e-06 | 1.958e-08 | 4.0156e-09 | 11 | 1
152 |  | sec 3 vs sec 4 reverse | 1.0041 | -1.0893e-05 | -2.7452e-05 | 6.3153e-06 | 5.4975e-09 | 5.4e-10 | [14 empty] | 1.0041 | -1.0893e-05 | -2.7452e-05 | 6.3153e-06 | 5.4975e-09 | 5.4e-10 | 13 | 1
153 |  | sex 3 vs sec 5 | 1.0052 | 3.1406e-06 | 7.8668e-06 | 5.4394e-06 | 5.994e-08 | 2.1855e-08 | 20 | 0.2 | 10 | 100.4074 | -0.2415 | 1.255 | -0.048104 | 0.24998 | [6 empty] | 1.0052 | 3.1406e-06 | 7.8668e-06 | 5.4394e-06 | 5.994e-08 | 2.1855e-08 | 11 | 1
154 |  | sex 3 vs sec 5 reverse | 1.003 | -7.9706e-06 | 1.7233e-05 | 6.7725e-06 | 4.3411e-08 | 7.8756e-08 | [14 empty] | 1.003 | -7.9706e-06 | 1.7233e-05 | 6.7725e-06 | 4.3411e-08 | 7.8756e-08 | 12 | 1
155 |  | sec 4 vs sec 5 | 1.007 | -1.2971e-07 | -2.9173e-06 | 6.7069e-06 | 3.986e-08 | 3.3624e-08 | 20 | 0.2 | 10 | 100.3968 | -0.24182 | -0.84362 | -0.048172 | -0.16806 | [6 empty] | 1.007 | -1.2971e-07 | -2.9173e-06 | 6.7069e-06 | 3.986e-08 | 3.3624e-08 | 10 | 1
156 |  | sec 4 vs sec 5 reverse | 1.0009 | -4.7066e-06 | -1.3955e-05 | 5.7884e-06 | 4.9786e-08 | 3.2951e-08 | [14 empty] | 1.0009 | -4.7066e-06 | -1.3955e-05 | 5.7884e-06 | 4.9786e-08 | 3.2951e-08 | 12 | 1
157
158 |  | Retest using 50A, after recrimping sections 2,3 & 5 and rewire section4, test to the washer just behind the series network.
159 | 2018-07-24 | Voltage drops,outer primary of Ta |  | measure across each section in turn, do not include any link | [6 empty] | section resistance | section g | Ik/Ips |  |  | Shunt V | Section mV | Shunt V
160 |  | section 1 | 0.50435 | 0.067841 | 0.13451 | 0.2 | 1.34 | 2.34 | 0.11111 | 1 | 0.0022031 | 453.9144 | 0.99422 |  |  | 0.50456 | 0.067841 | 0.50415
161 |  | section 2 | 0.5038 | 0.067694 | 0.13437 | 0.2 | 1.34 | 2.34 | 0.11097 | 2 | 0.0022026 | 454.0026 | 0.99441 |  |  | 0.50393 | 0.067694 | 0.50367
162 |  | section 3 | 0.50337 | 0.067276 | 0.13365 | 0.2 | 1.34 | 2.34 | 0.11025 | 3 | 0.0021903 | 456.5585 | 1 |  |  | 0.50355 | 0.067276 | 0.50318
163 |  | section 4 | 0.50306 | 0.066503 | 0.1322 | 0.2 | 1.34 | 2.34 | 0.1088 | 4 | 0.0021627 | 462.3771 | 1.0128 |  |  | 0.50311 | 0.066503 | 0.503
164 |  | section 5 | 0.5028 | 0.067217 | 0.13368 | 0.2 | 1.34 | 2.34 | 0.11028 | 5 | 0.0021934 | 455.9162 | 0.9986 |  |  | 0.50292 | 0.067217 | 0.50269
165
166 | 2018-07-24 | Magnetic coupling equality, outer primary of Ta | [6 empty] | CT ratio | Shunt | Com R | test current | Aux sec I |  | N(m2-m1)/Z1 |  |  | dm x dI | dm x dI | S/P error
167 | [12 empty] | Re uA | Im uA | Re ppm | Im ppm | Ipk/Ips | Re | Im | Re | Im
168 |  | sec 1 vs sec 2 | 0.97831 | -4.4126e-06 | -2.2878e-05 | 3.8915e-06 | 8.4703e-22 | 5.375e-08 | 20 | 0.2 | 10 | 97.8311 | -0.44126 | -2.2878 | -0.090209 | -0.4677 | [6 empty] | 0.97831 | -4.4126e-06 | -2.2878e-05 | 3.8915e-06 | 8.4703e-22 | 5.375e-08 | 13 | 1
169 |  | 27k across sec 1 | 0.97532 | -4.3525e-06 | -2.2524e-05 | 3.0489e-06 | 9.0898e-09 | 4.5568e-08 | 20 | 0.2 | 10 | 97.5323 | -0.43525 | -2.2524 | -0.089252 | -0.46188 | [6 empty] | 0.97532 | -4.3525e-06 | -2.2524e-05 | 3.0489e-06 | 9.0898e-09 | 4.5568e-08 | 13 | 1
170 |  | 27k across sec 2 | 0.97079 | -5.1654e-06 | -2.2693e-05 | 1.8897e-06 | 6.9952e-08 | 4.0343e-08 | 20 | 0.2 | 10 | 97.0786 | -0.51654 | -2.2693 | -0.10642 | -0.46752 | [6 empty] | 0.97079 | -5.1654e-06 | -2.2693e-05 | 1.8897e-06 | 6.9952e-08 | 4.0343e-08 | 13 | 1
171
172 | [11 empty] | Applied | Aux sec current, uA |  | N(m2-m1)/Z1 |  |  | dm x dI | dm x dI | S/P correction
173 | 2018-07-25 | Magnetic coupling equality, outer primary of Ta | [6 empty] | CT ratio | Shunt | Aux Sec res | current | Re | Im | Re ppm | Im ppm | Ipk/Ips |  |  | Re | Im
174 |  | sec 1 vs sec 2 | 1.004 | -6.0711e-06 | -1.8113e-05 | 2.6656e-06 | 3.0378e-08 | 2.8279e-08 | 20 | 0.2 | 10 | 100.393 | -0.6203 | -1.2469 | -0.12357 | -0.24841 | 0.0055878 | -0.0006905 | -0.001388 |  |  |  | 1.004 | -6.0711e-06 | -1.8113e-05 | 2.6656e-06 | 3.0378e-08 | 2.8279e-08 | 12 | 1
175 |  | sec 1 vs sec 2 reverse | 1.0038 | -6.3348e-06 | -6.8257e-06 | 3.4717e-06 | 1.3037e-08 | 1.4445e-08 | [14 empty] | 1.0038 | -6.3348e-06 | -6.8257e-06 | 3.4717e-06 | 1.3037e-08 | 1.4445e-08 | 11 | 1
176 |  | sec 1 vs sec3 | 0.99963 | -4.6672e-06 | -1.4566e-05 | 2.8548e-06 | 4.1828e-08 | 2.3954e-09 | 20 | 0.2 | 10 | 100.1578 | -0.57567 | -1.1037 | -0.11495 | -0.2204 | -1.0291e-05 | 1.183e-06 | 2.2682e-06 |  |  |  | 0.99963 | -4.6672e-06 | -1.4566e-05 | 2.8548e-06 | 4.1828e-08 | 2.3954e-09 | 12 | 1
177 |  | sec 1 vs sec3 reverse | 1.0035 | -6.8463e-06 | -7.5086e-06 | 3.0698e-06 | 3.6552e-09 | 2.3282e-08 | [14 empty] | 1.0035 | -6.8463e-06 | -7.5086e-06 | 3.0698e-06 | 3.6552e-09 | 2.3282e-08 | 11 | 1
178 |  | sec 1 vs sec 4 | 1.0032 | -3.2145e-06 | -8.2107e-06 | 3.305e-06 | 2.9159e-08 | 2.0031e-08 | 20 | 0.2 | 10 | 100.215 | -0.58984 | -1.1042 | -0.11771 | -0.22038 | -0.012755 | 0.0015015 | 0.0028109 |  |  |  | 1.0032 | -3.2145e-06 | -8.2107e-06 | 3.305e-06 | 2.9159e-08 | 2.0031e-08 | 11 | 1
179 |  | sec 1 vs sec 4 reverse | 1.0011 | -8.5823e-06 | -1.3874e-05 | 3.1657e-06 | 3.5547e-08 | 5.8948e-09 | [14 empty] | 1.0011 | -8.5823e-06 | -1.3874e-05 | 3.1657e-06 | 3.5547e-08 | 5.8948e-09 | 12 | 1
180 |  | sec 1 vs sec 5 | 1.0034 | -3.2116e-06 | -8.0265e-06 | 3.5086e-06 | 2.8234e-08 | 3.2865e-08 | 20 | 0.2 | 10 | 100.2048 | -0.37248 | -0.66737 | -0.074343 | -0.1332 | 0.0013965 | -0.00010382 | -0.00018601 | 0.00070832 | 0.0012391 | 2018 | 1.0034 | -3.2116e-06 | -8.0265e-06 | 3.5086e-06 | 2.8234e-08 | 3.2865e-08 | 11 | 1
181 |  | sec 1 vs sec 5 reverse | 1.0007 | -4.2379e-06 | -5.3209e-06 | 3.5483e-06 | 2.637e-08 | 2.0541e-08 | [11 empty] | -0.00012364 | 0.00011109 | 2013 | 1.0007 | -4.2379e-06 | -5.3209e-06 | 3.5483e-06 | 2.637e-08 | 2.0541e-08 | 11 | 1
182 |  | sec 2 vs sec 3 | 1.0013 | -8.1228e-07 | -3.9561e-06 | 3.2337e-06 | 2.3714e-08 | 2.3949e-08 | 20 | 0.2 | 10 | 100.2787 | -0.22709 | -0.46024 | -0.045292 | -0.091793 |  |  |  | -0.00030375 | -0.0015001 | 2000 | 1.0013 | -8.1228e-07 | -3.9561e-06 | 3.2337e-06 | 2.3714e-08 | 2.3949e-08 | 10 | 1
183 |  | sec 2 vs sec 3 reverse | 1.0043 | -3.7295e-06 | -5.2488e-06 | 3.5428e-06 | 3.0433e-08 | 3.2806e-08 | [14 empty] | 1.0043 | -3.7295e-06 | -5.2488e-06 | 3.5428e-06 | 3.0433e-08 | 3.2806e-08 | 11 | 1
184 |  | sec 2 vs sec 4 | 1.0019 | 5.3901e-09 | 1.3983e-06 | 3.0131e-06 | 4.5164e-08 | 2.9293e-08 | 20 | 0.2 | 10 | 100.0416 | -0.2456 | -0.382 | -0.0491 | -0.076369 | [6 empty] | 1.0019 | 5.3901e-09 | 1.3983e-06 | 3.0131e-06 | 4.5164e-08 | 2.9293e-08 | 9 | 1
185 |  | sec 2 vs sec 4 reverse | 0.99896 | -4.9174e-06 | -9.0383e-06 | 3.0217e-06 | 8.8888e-09 | 4.0795e-08 | [14 empty] | 0.99896 | -4.9174e-06 | -9.0383e-06 | 3.0217e-06 | 8.8888e-09 | 4.0795e-08 | 11 | 1
186 |  | sec 2 vs sec 5 | 1.0025 | -3.872e-06 | -9.7225e-07 | 3.2195e-06 | 1.7679e-08 | 2.4968e-08 | 20 | 0.2 | 10 | 100.2911 | -0.3899 | -0.42522 | -0.077753 | -0.084798 | [6 empty] | 1.0025 | -3.872e-06 | -9.7225e-07 | 3.2195e-06 | 1.7679e-08 | 2.4968e-08 | 10 | 1
187 |  | sec 2 vs sec 5 reverse | 1.0033 | -3.926e-06 | -7.5322e-06 | 3.6993e-06 | 1.4206e-08 | 4.4543e-08 | [14 empty] | 1.0033 | -3.926e-06 | -7.5322e-06 | 3.6993e-06 | 1.4206e-08 | 4.4543e-08 | 11 | 1
188 |  | sec 3 vs sec 4 | 1.0083 | -4.4683e-07 | -7.3371e-07 | 4.4935e-06 | 2.9315e-08 | 3.0092e-08 | 20 | 0.2 | 10 | 100.5526 | -0.205 | -0.41538 | -0.040776 | -0.082619 | [6 empty] | 1.0083 | -4.4683e-07 | -7.3371e-07 | 4.4935e-06 | 2.9315e-08 | 3.0092e-08 | 8 | 1
189 |  | sec 3 vs sec 4 reverse | 1.0028 | -3.6533e-06 | -7.5739e-06 | 3.7241e-06 | 2.6274e-08 | 2.3607e-08 | [14 empty] | 1.0028 | -3.6533e-06 | -7.5739e-06 | 3.7241e-06 | 2.6274e-08 | 2.3607e-08 | 11 | 1
190 |  | sex 3 vs sec 5 | 0.99632 | -3.0295e-06 | -4.9528e-06 | 3.2738e-06 | 3.3533e-08 | 2.4071e-08 | 20 | 0.2 | 10 | 100.0435 | -0.26815 | -0.43403 | -0.053607 | -0.086769 | [6 empty] | 0.99632 | -3.0295e-06 | -4.9528e-06 | 3.2738e-06 | 3.3533e-08 | 2.4071e-08 | 11 | 1
191 |  | sex 3 vs sec 5 reverse | 1.0045 | -2.3335e-06 | -3.7278e-06 | 2.8622e-06 | 2.2797e-08 | 1.7735e-08 | [14 empty] | 1.0045 | -2.3335e-06 | -3.7278e-06 | 2.8622e-06 | 2.2797e-08 | 1.7735e-08 | 10 | 1
192 |  | sec 4 vs sec 5 | 1.0014 | -4.9991e-06 | -1.1073e-05 | 2.6564e-06 | 4.9684e-08 | 4.1087e-09 | 20 | 0.2 | 10 | 100.1261 | -0.32454 | -0.63896 | -0.064825 | -0.12763 | [6 empty] | 1.0014 | -4.9991e-06 | -1.1073e-05 | 2.6564e-06 | 4.9684e-08 | 4.1087e-09 | 12 | 1
193 |  | sec 4 vs sec 5 reverse | 1.0012 | -1.4916e-06 | -1.7067e-06 | 3.0532e-06 | 2.2926e-08 | 9.8307e-09 | [14 empty] | 1.0012 | -1.4916e-06 | -1.7067e-06 | 3.0532e-06 | 2.2926e-08 | 9.8307e-09 | 10 | 1
194
195
196 | [8 empty] | CT ratio | Shunt | Com R | test current | Aux sec I |  | N(m2-m1)/Z1 |  |  | dm x dI | dm x dI | S/P error
197 | 2018-07-25 | Magnetic coupling equality, outer primary of Tb | [10 empty] | Re uA | Im uA | Re ppm | Im ppm | 1-Ipk/Ips | Re | Im | Re | Im
198 |  | sec 1 vs sec 2 | 1.0222 | -1.3994e-06 | -1.1383e-05 | 5.3089e-06 | 1.8524e-09 | 2.4266e-08 | 20 | 0.2 | 10 | 102.2158 | -0.13994 | -1.1383 | -0.027381 | -0.22273 | -1.2899e-05 | 3.5318e-07 | 2.8729e-06 | 3.5318e-07 | 2.8729e-06 |  | 1.0222 | -1.3994e-06 | -1.1383e-05 | 5.3089e-06 | 1.8524e-09 | 2.4266e-08 | 12 | 1
199 |  | 27k across sec 1 | 1.0178 | -1.6647e-08 | -1.0823e-05 | 3.7728e-06 | 7.9829e-09 | 1.4954e-08 | 20 | 0.2 | 10 | 101.7828 | -0.0016647 | -1.0823 | [5 empty] | -0.0002282 | -0.00091269 |  | 1.0178 | -1.6647e-08 | -1.0823e-05 | 3.7728e-06 | 7.9829e-09 | 1.4954e-08 | 11 | 1
200 |  | 27k across sec 2 | 1.0138 | -2.6598e-06 | -1.0701e-05 | 3.995e-06 | 1.2192e-08 | 1.4742e-08 | 20 | 0.2 | 10 | 101.3756 | -0.26598 | -1.0701 | [5 empty] | -0.00048473 | -0.00018268 |  | 1.0138 | -2.6598e-06 | -1.0701e-05 | 3.995e-06 | 1.2192e-08 | 1.4742e-08 | 11 | 1
201
202 | [6 empty] | sec internal | total | pri only |  | V drop | cond S | relative | [6 empty] | Cumulative
203 | Repeat | Volt drops, outer primary of Tb |  |  | sec re 1 V | shunt | burden V | sec V | V drop | section | ea section | ea section | current |  |  | Shunt V | Section mV | Shunt V |  | section V
204 |  | Volt drop section 1 | 0.50044 | 0.00010785 | 0.00021551 | 0.2 | 1.585 | 2.585 | -0.12903 | 1 | -0.12903 | -7.7499 | 0.99999 |  |  | 0.50055 | 0.10785 | 0.50032 |  | 0.10785
205 |  | Volt drop sections 1+2 | 0.49993 | 0.00021714 | 0.00043435 | 0.2 | 1.585 | 2.585 | -0.25807 | 2 | -0.12903 | -7.7501 | 1 |  |  | 0.50011 | 0.1093 | 0.49975 |  | 0.21714
206
207
208 | Repeat | Voltage drops, middle primary of Ta |  |  |  |  | sec internal | total | pri only |  | V drop | cond S | relative | [6 empty] | Cumulative
209 |  |  |  |  | sec re 1 V | shunt | burden V | sec V | V drop | section | ea section | ea section | current |  |  | Shunt V | Section V | Shunt V |  | section V
210 |  | Volt drop section 1 | 1.0067 | 0.3608 | 0.35839 | 0.2 | 1.34 | 2.34 | 0.24139 | 1 | 0.24139 | 4.1427 | 1.0558 |  |  | 1.0077 | 0.3608 | 1.0058 |  | 0.3608
211 |  | Volt drop sections 1+2 | 1.0047 | 0.75805 | 0.75449 | 0.2 | 1.34 | 2.34 | 0.52049 | 2 | 0.2791 | 3.583 | 0.91317 |  |  | 1.0052 | 0.39724 | 1.0043 |  | 0.75805
212 |  | Volt drop sections 1+2+3 | 1.0033 | 1.1427 | 1.139 | 0.2 | 1.34 | 2.34 | 0.78801 | 3 | 0.26752 | 3.738 | 0.95269 |  |  | 1.0039 | 0.38467 | 1.0026 |  | 1.1427
213 |  | Volt drop sections 1+2+3+4 | 1.0018 | 1.4951 | 1.4924 | 0.2 | 1.34 | 2.34 | 1.0244 | 4 | 0.23636 | 4.2309 | 1.0783 |  |  | 1.0023 | 0.35237 | 1.0013 |  | 1.4951
214
215 | Repeat | Magnetic coupling equality, middle primary of Ta | [6 empty] | CT ratio | Shunt | Com R | test current | Aux sec I |  | N(m2-m1)/Z1 |  |  | dm x dI | dm x dI | S/P error
216 |  | This was read with Ta and Tb LP terminals in series. Phase on CH set to the 0.2 ohm shunt. | [10 empty] | Re uA | Im uA | Re ppm | Im ppm | 1-Ipk/Ips | Re | Im | Re | Im
217 |  | sec 1 vs sec 2 | 1.0012 | 3.714e-07 | -5.032e-07 | 8.4951e-06 | 1.023e-08 | 1.9541e-08 | 4 | 0.2 | 10 | 20.0239 | 0.03714 | -0.05032 | 0.0074191 | -0.010052 | [6 empty] | 1.0012 | 3.714e-07 | -5.032e-07 | 8.4951e-06 | 1.023e-08 | 1.9541e-08 | 7 | 1
218 |  | sec 1 vs 2: 100k across 1 | 1.0054 | 1.3309e-06 | -2.3615e-07 | 2.8281e-06 | 1.3342e-08 | 1.8753e-08 | 4 | 0.2 | 10 | 20.1083 | 0.13309 | -0.023615 | 0.026475 | -0.0046976 | [6 empty] | 1.0054 | 1.3309e-06 | -2.3615e-07 | 2.8281e-06 | 1.3342e-08 | 1.8753e-08 | 9 | 1
219
220 |  | sec 1 vs sec 2 | 1.0034 | 3.909e-07 | -4.5404e-07 | 1.8685e-06 | 1.7579e-08 | 1.4725e-08 | 4 | 0.2 | 10 | 20.0679 | 0.03909 | -0.045404 | 0.0077916 | -0.0090501 | 0.086826 | 0.00067651 | -0.00078578 |  |  |  | 1.0034 | 3.909e-07 | -4.5404e-07 | 1.8685e-06 | 1.7579e-08 | 1.4725e-08 | 8 | 1
221 |  | sec 1 vs sec3 | 1.0019 | 7.2293e-07 | 8.5118e-07 | 2.9628e-06 | 1.4349e-08 | 1.3286e-08 | 4 | 0.2 | 10 | 20.0383 | 0.072293 | 0.085118 | 0.014431 | 0.016991 | 0.047308 | 0.00068269 | 0.00080381 |  |  |  | 1.0019 | 7.2293e-07 | 8.5118e-07 | 2.9628e-06 | 1.4349e-08 | 1.3286e-08 | 9 | 1
222 |  | sec 1 vs sec 4 | 1.0036 | -2.0288e-06 | -5.1614e-06 | 3.481e-06 | 2.3063e-08 | 1.7932e-08 | 4 | 0.2 | 10 | 20.0728 | -0.20288 | -0.51614 | -0.040428 | -0.10285 | -0.078301 | 0.0031656 | 0.0080535 | 0.0045248 | 0.0080715 |  | 1.0036 | -2.0288e-06 | -5.1614e-06 | 3.481e-06 | 2.3063e-08 | 1.7932e-08 | 11 | 1
223 |  | sec 2 vs sec3 | 1.0033 | -1.5873e-06 | -2.6233e-06 | 3.4777e-06 | 1.4069e-08 | 1.2137e-08 | 4 | 0.2 | 10 | 20.0657 | -0.15873 | -0.26233 | -0.031642 | -0.052294 |  |  |  | 0.001573 | 0.00069226 |  | 1.0033 | -1.5873e-06 | -2.6233e-06 | 3.4777e-06 | 1.4069e-08 | 1.2137e-08 | 10 | 1
224 |  | sec 2 vs sec 4 | 1.0007 | -4.4394e-06 | -8.8644e-06 | 1.9042e-05 | 1.6576e-08 | 1.0756e-08 | 4 | 0.2 | 10 | 20.0135 | -0.44394 | -0.88644 | -0.088729 | -0.17717 |  |  |  | 0.0048877 | -0.0034944 |  | 1.0007 | -4.4394e-06 | -8.8644e-06 | 1.9042e-05 | 1.6576e-08 | 1.0756e-08 | 11 | 1
225 |  | sec 3 vs sec 4 | 0.99882 | -4.8497e-06 | -1.0379e-05 | 4.4154e-06 | 2.0809e-09 | 3.6103e-09 | 4 | 0.2 | 10 | 19.9765 | -0.48497 | -1.0379 | -0.097107 | -0.20783 | [6 empty] | 0.99882 | -4.8497e-06 | -1.0379e-05 | 4.4154e-06 | 2.0809e-09 | 3.6103e-09 | 12 | 1
226
227
| from above | sec 1 vs sec 2 | 1.0039 | -6.203e-06 | -1.2469e-05 | 3.0687e-06 | 2.1708e-08 | 2.1362e-08 | 20 | 0.2 | 10
|  | sec 1 vs sec3 | 1.0016 | -5.7567e-06 | -1.1037e-05 | 2.9623e-06 | 2.2742e-08 | 1.2839e-08 | 20 | 0.2 | 10
|  | sec 1 vs sec 4 | 1.0021 | -5.8984e-06 | -1.1042e-05 | 3.2353e-06 | 3.2353e-08 | 1.2963e-08 | 20 | 0.2 | 10
|  | sec 1 vs sec 5 | 1.002 | -3.7248e-06 | -6.6737e-06 | 3.5284e-06 | 2.7302e-08 | 2.6703e-08 | 20 | 0.2 | 10

## PythonResults
| e1.real | e2.real | e3.real | e4.real | e5a.real | e5b.real
125 | 2.1318e-07 | 5.4168e-08 | 1.3799e-07 | 1.593e-08 | 1.7293e-07 | 1.663e-07
120 | 2.1686e-07 | 7.6805e-08 | 1.0671e-07 | 1.9587e-08 | 1.8354e-07 | 1.7221e-07
100 | 2.379e-07 | 5.9966e-08 | 1.0918e-07 | 3.1378e-08 | 1.9133e-07 | 1.8116e-07
60 | 3.196e-07 | 6.5014e-08 | 1.5199e-07 | 4.9731e-08 | 2.1748e-07 | 2.1384e-07
40 | 3.7847e-07 | 6.4653e-08 | 1.8548e-07 | 5.5829e-08 | 2.535e-07 | 2.4482e-07
20 | 4.5019e-07 | 7.3141e-08 | 2.1984e-07 | 6.7545e-08 | 2.9032e-07 | 2.7574e-07
10 | 4.9688e-07 | 8.4832e-08 | 2.4561e-07 | 7.3678e-08 | 3.174e-07 | 3.0294e-07
5 | 5.4764e-07 | 8.4647e-08 | 2.6103e-07 | 7.7827e-08 | 3.351e-07 | 3.193e-07
1 | 5.8247e-07 | 1.0293e-07 | 2.8463e-07 | 8.3829e-08 | 3.4832e-07 | 3.571e-07
| e1.imag | e2.imag | e3.imag | e4.imag | e5a.imag | e5b.imag
125 | 1.0083e-07 | 3.8801e-08 | 1.1146e-07 | 1.7449e-08 | 7.7527e-08 | 8.5833e-08
120 | 1.019e-07 | 3.6709e-08 | 1.1023e-07 | 1.7716e-08 | 8.4741e-08 | 8.5434e-08
100 | 1.0402e-07 | 3.8893e-08 | 1.1544e-07 | 1.7694e-08 | 8.6556e-08 | 8.6842e-08
60 | 9.2616e-08 | 3.0893e-08 | 1.2676e-07 | 1.4157e-08 | 1.002e-07 | 9.5616e-08
40 | 8.1632e-08 | 4.0083e-08 | 1.2351e-07 | 1.8117e-08 | 9.5909e-08 | 8.9902e-08
20 | 7.5248e-08 | 4.6122e-08 | 1.249e-07 | 2.2282e-08 | 9.3312e-08 | 8.7485e-08
10 | 6.7684e-08 | 4.1995e-08 | 1.2236e-07 | 2.3831e-08 | 8.6754e-08 | 8.6138e-08
5 | 6.2711e-08 | 4.7726e-08 | 1.2706e-07 | 2.8571e-08 | 8.9642e-08 | 9.6145e-08
1 | 7.3157e-08 | 6.7979e-08 | 1.1749e-07 | 2.0923e-08 | 9.799e-08 | 1.0102e-07
| 1a | 1b | 2a | 3a | 2b
real | 6.476170636821842e-10 | 1.968546386339233e-08 | -1.0622402663419591e-07 | -6.132813765550067e-08 | -7.094267261374557e-11
imag | -2.0406753719131177e-09 | -4.772868736570222e-09 | -1.8427993259160942e-07 | -1.0729080164442446e-07 | -5.770797066207241e-10
| real | imag
capa s - p | ureal(-1.08068314064672351343455e-08,2.16136628129344694415105e-09,5) | ureal(1.495084437858827370981e-07,2.9901688757176547419619e-08,5)
capb s - p | ureal(-7.937033481081537829866893778e-13,1.58740669621630786889596636e-13,5) | ureal(8.46164652028692852541406e-09,1.69232930405738591187797e-09,5)
P1as
Excitation | Real | Imag | U real | U imag | k real | k imag
125 | 2.1318e-07 | 1.0083e-07 | 3.7235e-08 | 2.3912e-08 | 2.107 | 2.1122
120 | 2.1686e-07 | 1.019e-07 | 3.7257e-08 | 2.3939e-08 | 2.1066 | 2.1113
100 | 2.379e-07 | 1.0402e-07 | 3.75e-08 | 2.403e-08 | 2.1022 | 2.1082
60 | 3.196e-07 | 9.2616e-08 | 3.8777e-08 | 2.45e-08 | 2.0846 | 2.0959
40 | 3.7847e-07 | 8.1632e-08 | 4.0259e-08 | 2.548e-08 | 2.0713 | 2.0842
20 | 4.5019e-07 | 7.5248e-08 | 4.5351e-08 | 3.1879e-08 | 2.0525 | 2.1035
10 | 4.9688e-07 | 6.7684e-08 | 6.0134e-08 | 5.2488e-08 | 2.11 | 2.2653
5 | 5.4764e-07 | 6.2711e-08 | 1.0401e-07 | 1.0192e-07 | 2.2856 | 2.4338
1 | 5.8247e-07 | 7.3157e-08 | 4.9568e-07 | 4.9618e-07 | 2.4866 | 2.5221
P1ap
Excitation | Real | Imag | U real | U imag | k real | k imag
125 | 2.0302e-07 | 2.4829e-07 | 5.1607e-08 | 7.8642e-08 | 2.0271 | 2.2635
120 | 2.067e-07 | 2.4937e-07 | 5.168e-08 | 7.8655e-08 | 2.0269 | 2.2629
100 | 2.2775e-07 | 2.5149e-07 | 5.2212e-08 | 7.8687e-08 | 2.0249 | 2.2616
60 | 3.0944e-07 | 2.4008e-07 | 5.4825e-08 | 7.8693e-08 | 2.0189 | 2.2608
40 | 3.6831e-07 | 2.291e-07 | 5.7384e-08 | 7.879e-08 | 2.0165 | 2.2566
20 | 4.4003e-07 | 2.2272e-07 | 6.3017e-08 | 7.9981e-08 | 2.0137 | 2.2214
10 | 4.8672e-07 | 2.1515e-07 | 7.4652e-08 | 8.6268e-08 | 2.0337 | 2.1566
5 | 5.3748e-07 | 2.1018e-07 | 1.1109e-07 | 1.1564e-07 | 2.1544 | 2.1796
1 | 5.7231e-07 | 2.2062e-07 | 4.9559e-07 | 4.9573e-07 | 2.4667 | 2.4859
P2ap
Excitation | Real | Imag | U real | U imag | k real | k imag
125 | 3.0653e-07 | 3.2464e-07 | 8.0239e-08 | 8.4816e-08 | 1.9872 | 2.1286
120 | 2.5629e-07 | 3.2658e-07 | 8.0373e-08 | 8.4848e-08 | 1.987 | 2.1282
100 | 2.9664e-07 | 3.3173e-07 | 8.0651e-08 | 8.4939e-08 | 1.9868 | 2.1271
60 | 4.161e-07 | 3.3964e-07 | 8.2692e-08 | 8.5265e-08 | 1.9853 | 2.1229
40 | 5.0882e-07 | 3.1621e-07 | 8.4867e-08 | 8.5851e-08 | 1.9844 | 2.1162
20 | 6.0641e-07 | 3.0518e-07 | 9.1932e-08 | 9.0263e-08 | 1.9826 | 2.0818
10 | 6.6718e-07 | 2.9921e-07 | 1.1376e-07 | 1.1007e-07 | 1.9909 | 2.0347
5 | 7.3355e-07 | 2.932e-07 | 1.7001e-07 | 1.6565e-07 | 2.0394 | 2.0595
1 | 7.7369e-07 | 2.7382e-07 | 7.4302e-07 | 7.4115e-07 | 2.1143 | 2.1168
P3as
Excitation | Real | Imag | U real | U imag | k real | k imag
125 | 4.6353e-07 | 3.8472e-07 | 9.3281e-08 | 8.9344e-08 | 1.98 | 2.0851
120 | 4.2024e-07 | 3.9361e-07 | 9.336e-08 | 8.9366e-08 | 1.98 | 2.0849
100 | 4.566e-07 | 4.0059e-07 | 9.3582e-08 | 8.9503e-08 | 1.9799 | 2.0838
60 | 5.8385e-07 | 4.2568e-07 | 9.5641e-08 | 9.012e-08 | 1.9788 | 2.0791
40 | 7.0649e-07 | 3.9401e-07 | 9.8192e-08 | 9.1327e-08 | 1.9779 | 2.0711
20 | 8.2919e-07 | 3.7621e-07 | 1.0736e-07 | 9.866e-08 | 1.9761 | 2.0398
10 | 9.109e-07 | 3.6213e-07 | 1.3555e-07 | 1.2691e-07 | 1.9832 | 2.0098
5 | 9.9082e-07 | 3.5427e-07 | 2.1196e-07 | 2.0455e-07 | 2.0115 | 2.0237
1 | 1.0382e-06 | 3.5089e-07 | 9.3488e-07 | 9.3395e-07 | 2.0495 | 2.0499
P3as
Excitation | Real | Imag | U real | U imag | k real | k imag
125 | 3.3484e-07 | 2.9901e-07 | 8.0083e-08 | 8.4815e-08 | 1.9874 | 2.1286
120 | 3.2179e-07 | 3.0178e-07 | 8.009e-08 | 8.4841e-08 | 1.9874 | 2.1283
100 | 3.6862e-07 | 3.0313e-07 | 8.0606e-08 | 8.4929e-08 | 1.9869 | 2.1272
60 | 4.7795e-07 | 3.085e-07 | 8.2545e-08 | 8.5202e-08 | 1.9855 | 2.1236
40 | 5.6816e-07 | 2.8261e-07 | 8.4935e-08 | 8.5902e-08 | 1.9843 | 2.1156
20 | 6.6231e-07 | 2.6777e-07 | 9.2987e-08 | 9.0248e-08 | 1.9816 | 2.0819
10 | 7.245e-07 | 2.6298e-07 | 1.1438e-07 | 1.1022e-07 | 1.9902 | 2.0343
5 | 7.9182e-07 | 2.6229e-07 | 1.7266e-07 | 1.6735e-07 | 2.035 | 2.0561
1 | 8.4616e-07 | 2.5735e-07 | 7.5979e-07 | 7.6023e-07 | 2.0985 | 2.0981
P1bs
Excitation | Real | Imag | U real | U imag | k real | k imag
125 | 1.4885e-07 | 2.0949e-07 | 6.1542e-08 | 8.0444e-08 | 2.0069 | 2.2052
120 | 1.299e-07 | 2.1266e-07 | 6.159e-08 | 8.0464e-08 | 2.0068 | 2.2047
100 | 1.6778e-07 | 2.126e-07 | 6.222e-08 | 8.0522e-08 | 2.0052 | 2.2032
60 | 2.4443e-07 | 2.0919e-07 | 6.4494e-08 | 8.064e-08 | 2.0021 | 2.1999
40 | 3.0365e-07 | 1.8902e-07 | 6.7072e-08 | 8.099e-08 | 1.9999 | 2.1915
20 | 3.6689e-07 | 1.7659e-07 | 7.3919e-08 | 8.3773e-08 | 1.9969 | 2.1467
10 | 4.0188e-07 | 1.7316e-07 | 9.3335e-08 | 9.8035e-08 | 2.0078 | 2.0736
5 | 4.5283e-07 | 1.6245e-07 | 1.4208e-07 | 1.4225e-07 | 2.0822 | 2.1035
1 | 4.6937e-07 | 1.5265e-07 | 6.2922e-07 | 6.2932e-07 | 2.2022 | 2.202
P1bp
Excitation | Real | Imag | U real | U imag | k real | k imag
125 | 1.6854e-07 | 2.1318e-07 | 7.2703e-08 | 8.264e-08 | 1.9934 | 2.1599
120 | 1.4958e-07 | 2.1635e-07 | 7.2744e-08 | 8.2662e-08 | 1.9934 | 2.1595
100 | 1.8746e-07 | 2.1629e-07 | 7.3282e-08 | 8.2725e-08 | 1.9926 | 2.1584
60 | 2.6411e-07 | 2.1288e-07 | 7.5222e-08 | 8.2857e-08 | 1.9909 | 2.156
40 | 3.2334e-07 | 1.9271e-07 | 7.7434e-08 | 8.3236e-08 | 1.9897 | 2.1499
20 | 3.8657e-07 | 1.8028e-07 | 8.3404e-08 | 8.6139e-08 | 1.988 | 2.1167
10 | 4.2157e-07 | 1.7685e-07 | 1.0085e-07 | 1.003e-07 | 1.9969 | 2.061
5 | 4.7252e-07 | 1.6614e-07 | 1.4661e-07 | 1.4373e-07 | 2.0639 | 2.0951
1 | 4.8906e-07 | 1.5633e-07 | 6.2999e-07 | 6.2958e-07 | 2.1997 | 2.2011
P2bs
Excitation | Real | Imag | U real | U imag | k real | k imag
125 | 2.906e-07 | 3.0719e-07 | 8.6963e-08 | 8.7055e-08 | 1.9831 | 2.1044
120 | 2.367e-07 | 3.0886e-07 | 8.7072e-08 | 8.7092e-08 | 1.983 | 2.104
100 | 2.6526e-07 | 3.1404e-07 | 8.7356e-08 | 8.7206e-08 | 1.9828 | 2.1029
60 | 3.6637e-07 | 3.2548e-07 | 8.9394e-08 | 8.7682e-08 | 1.9816 | 2.0983
40 | 4.5299e-07 | 2.981e-07 | 9.1697e-08 | 8.8534e-08 | 1.9807 | 2.091
20 | 5.3887e-07 | 2.829e-07 | 9.9768e-08 | 9.4389e-08 | 1.979 | 2.0578
10 | 5.935e-07 | 2.7538e-07 | 1.2512e-07 | 1.1871e-07 | 1.9865 | 2.0199
5 | 6.5572e-07 | 2.6463e-07 | 1.9312e-07 | 1.8676e-07 | 2.0209 | 2.0361
1 | 6.8986e-07 | 2.529e-07 | 8.4552e-07 | 8.4451e-07 | 2.072 | 2.0728
P3ap
Excitation | Real | Imag | U real | U imag | k real | k imag
125 | 2.7351e-07 | 1.9172e-07 | 1.0318e-07 | 1.4717e-07 | 2.0375 | 2.2023
120 | 2.6047e-07 | 1.9449e-07 | 1.0319e-07 | 1.4718e-07 | 2.0375 | 2.2021
100 | 3.0729e-07 | 1.9584e-07 | 1.0356e-07 | 1.4722e-07 | 2.0362 | 2.2015
60 | 4.1662e-07 | 2.012e-07 | 1.0498e-07 | 1.4736e-07 | 2.0319 | 2.1997
40 | 5.0683e-07 | 1.7532e-07 | 1.0675e-07 | 1.4769e-07 | 2.0272 | 2.1951
20 | 6.0098e-07 | 1.6048e-07 | 1.1293e-07 | 1.4976e-07 | 2.0144 | 2.1714
10 | 6.6318e-07 | 1.5569e-07 | 1.3043e-07 | 1.6043e-07 | 2.0022 | 2.1041
5 | 7.3049e-07 | 1.55e-07 | 1.8269e-07 | 2.0004e-07 | 2.0249 | 2.0521
1 | 7.8483e-07 | 1.5006e-07 | 7.6176e-07 | 7.6632e-07 | 2.0964 | 2.0921
P2as
Excitation | Real | Imag | U real | U imag | k real | k imag
125 | 4.1275e-07 | 5.0892e-07 | 1.4509e-07 | 2.3762e-07 | 2.1743 | 2.3669
120 | 3.6251e-07 | 5.1086e-07 | 1.4514e-07 | 2.3762e-07 | 2.1737 | 2.3668
100 | 4.0287e-07 | 5.1601e-07 | 1.4525e-07 | 2.3763e-07 | 2.1725 | 2.3663
60 | 5.2233e-07 | 5.2392e-07 | 1.4608e-07 | 2.3768e-07 | 2.1639 | 2.3645
40 | 6.1505e-07 | 5.0049e-07 | 1.4699e-07 | 2.3775e-07 | 2.1553 | 2.3614
20 | 7.1263e-07 | 4.8946e-07 | 1.5019e-07 | 2.3835e-07 | 2.13 | 2.3406
10 | 7.734e-07 | 4.8349e-07 | 1.6205e-07 | 2.422e-07 | 2.077 | 2.2668
5 | 8.3977e-07 | 4.7748e-07 | 2.0179e-07 | 2.6249e-07 | 2.0413 | 2.1461
1 | 8.7991e-07 | 4.581e-07 | 7.4896e-07 | 7.5962e-07 | 2.1073 | 2.0981
P2asp
Excitation | Real | Imag | U real | U imag | k real | k imag
125 | 3.1343e-07 | 3.5079e-07 | 1.7407e-07 | 2.8212e-07 | 2.1101 | 2.1799
120 | 2.6319e-07 | 3.5273e-07 | 1.7412e-07 | 2.8213e-07 | 2.1099 | 2.1798
100 | 3.0355e-07 | 3.5788e-07 | 1.7422e-07 | 2.8215e-07 | 2.1093 | 2.1797
60 | 4.2301e-07 | 3.6579e-07 | 1.75e-07 | 2.8224e-07 | 2.1054 | 2.1791
40 | 5.1573e-07 | 3.4237e-07 | 1.7586e-07 | 2.8239e-07 | 2.1014 | 2.1782
20 | 6.1331e-07 | 3.3133e-07 | 1.7878e-07 | 2.835e-07 | 2.089 | 2.1717
10 | 6.7408e-07 | 3.2536e-07 | 1.8916e-07 | 2.887e-07 | 2.0591 | 2.1466
5 | 7.4045e-07 | 3.1936e-07 | 2.2364e-07 | 3.0814e-07 | 2.0311 | 2.0951
1 | 7.8059e-07 | 2.9997e-07 | 7.5392e-07 | 7.7404e-07 | 2.1018 | 2.0855
P1asp
Excitation | Real | Imag | U real | U imag | k real | k imag
125 | 2.0461e-07 | 2.2213e-07 | 5.1545e-08 | 6.5132e-08 | 2.0275 | 2.1895
120 | 2.0829e-07 | 2.2321e-07 | 5.1618e-08 | 6.5152e-08 | 2.0273 | 2.189
100 | 2.2933e-07 | 2.2533e-07 | 5.2151e-08 | 6.5203e-08 | 2.0253 | 2.1877
60 | 3.1103e-07 | 2.1392e-07 | 5.4766e-08 | 6.5212e-08 | 2.0192 | 2.1867
40 | 3.6989e-07 | 2.0294e-07 | 5.7328e-08 | 6.5362e-08 | 2.0168 | 2.1825
20 | 4.4162e-07 | 1.9656e-07 | 6.2966e-08 | 6.7138e-08 | 2.0139 | 2.151
10 | 4.883e-07 | 1.8899e-07 | 7.461e-08 | 7.5609e-08 | 2.0339 | 2.115
5 | 5.3907e-07 | 1.8402e-07 | 1.1107e-07 | 1.0996e-07 | 2.1547 | 2.2025
1 | 5.739e-07 | 1.9446e-07 | 4.9559e-07 | 4.958e-07 | 2.4668 | 2.4963
P2bp
Excitation | Real | Imag | U real | U imag | k real | k imag
125 | 2.9053e-07 | 3.0661e-07 | 2.1408e-07 | 2.1242e-07 | 1.9606 | 1.9628
120 | 2.3663e-07 | 3.0829e-07 | 2.1412e-07 | 2.1244e-07 | 1.9606 | 1.9628
100 | 2.6519e-07 | 3.1346e-07 | 2.1424e-07 | 2.1249e-07 | 1.9606 | 1.9628
60 | 3.663e-07 | 3.249e-07 | 2.1508e-07 | 2.1273e-07 | 1.9606 | 1.9628
40 | 4.5292e-07 | 2.9752e-07 | 2.1605e-07 | 2.1315e-07 | 1.9606 | 1.9628
20 | 5.388e-07 | 2.8232e-07 | 2.1958e-07 | 2.1594e-07 | 1.9607 | 1.9628
10 | 5.9343e-07 | 2.748e-07 | 2.3188e-07 | 2.2778e-07 | 1.9621 | 1.9638
5 | 6.5565e-07 | 2.6406e-07 | 2.7298e-07 | 2.6806e-07 | 1.9735 | 1.9754
1 | 6.8979e-07 | 2.5232e-07 | 8.6511e-07 | 8.6409e-07 | 2.0591 | 2.0597

## Tables
Table 1
Kn = | 1 |  | 5:5 |  | Configuration: | P1as
Operating Conditions |  | Error |  | Expanded Uncertainty |  | Coverage Factor
Burden (Ω) | Primary current (% of rated) | Magnitude (µA/A) | Phase (µrad) | Magnitude (µA/A) | Phase (µrad) | Magnitude | Phase
0.2 | 125 | 0.21318 | 0.10083 | 0.037235 | 0.023912 | 2.107 | 2.1122
0.2 | 120 | 0.21686 | 0.1019 | 0.037257 | 0.023939 | 2.1066 | 2.1113
0.2 | 100 | 0.2379 | 0.10402 | 0.0375 | 0.02403 | 2.1022 | 2.1082
0.2 | 60 | 0.3196 | 0.092616 | 0.038777 | 0.0245 | 2.0846 | 2.0959
0.2 | 40 | 0.37847 | 0.081632 | 0.040259 | 0.02548 | 2.0713 | 2.0842
0.2 | 20 | 0.45019 | 0.075248 | 0.045351 | 0.031879 | 2.0525 | 2.1035
0.2 | 10 | 0.49688 | 0.067684 | 0.060134 | 0.052488 | 2.11 | 2.2653
0.2 | 5 | 0.54764 | 0.062711 | 0.10401 | 0.10192 | 2.2856 | 2.4338
0.2 | 1 | 0.58247 | 0.073157 | 0.49568 | 0.49618 | 2.4866 | 2.5221
Table 2
Kn = | 4 |  | 20:5 |  | Configuration: | P1ap
Operating Conditions |  | Error |  | Expanded Uncertainty |  | Coverage Factor
Burden (Ω) | Primary current (% of rated) | Magnitude (µA/A) | Phase (µrad) | Magnitude (µA/A) | Phase (µrad) | Magnitude | Phase
0.2 | 125 | 0.20302 | 0.24829 | 0.051607 | 0.078642 | 2.0271 | 2.2635
0.2 | 120 | 0.2067 | 0.24937 | 0.05168 | 0.078655 | 2.0269 | 2.2629
0.2 | 100 | 0.22775 | 0.25149 | 0.052212 | 0.078687 | 2.0249 | 2.2616
0.2 | 60 | 0.30944 | 0.24008 | 0.054825 | 0.078693 | 2.0189 | 2.2608
0.2 | 40 | 0.36831 | 0.2291 | 0.057384 | 0.07879 | 2.0165 | 2.2566
0.2 | 20 | 0.44003 | 0.22272 | 0.063017 | 0.079981 | 2.0137 | 2.2214
0.2 | 10 | 0.48672 | 0.21515 | 0.074652 | 0.086268 | 2.0337 | 2.1566
0.2 | 5 | 0.53748 | 0.21018 | 0.11109 | 0.11564 | 2.1544 | 2.1796
0.2 | 1 | 0.57231 | 0.22062 | 0.49559 | 0.49573 | 2.4667 | 2.4859
Table 3
Kn = | 20 |  | 100:5 |  | Configuration: | P2ap
Operating Conditions |  | Error |  | Expanded Uncertainty |  | Coverage Factor
Burden (Ω) | Primary current (% of rated) | Magnitude (µA/A) | Phase (µrad) | Magnitude (µA/A) | Phase (µrad) | Magnitude | Phase
0.2 | 125 | 0.30653 | 0.32464 | 0.080239 | 0.084816 | 1.9872 | 2.1286
0.2 | 120 | 0.25629 | 0.32658 | 0.080373 | 0.084848 | 1.987 | 2.1282
0.2 | 100 | 0.29664 | 0.33173 | 0.080651 | 0.084939 | 1.9868 | 2.1271
0.2 | 60 | 0.4161 | 0.33964 | 0.082692 | 0.085265 | 1.9853 | 2.1229
0.2 | 40 | 0.50882 | 0.31621 | 0.084867 | 0.085851 | 1.9844 | 2.1162
0.2 | 20 | 0.60641 | 0.30518 | 0.091932 | 0.090263 | 1.9826 | 2.0818
0.2 | 10 | 0.66718 | 0.29921 | 0.11376 | 0.11007 | 1.9909 | 2.0347
0.2 | 5 | 0.73355 | 0.2932 | 0.17001 | 0.16565 | 2.0394 | 2.0595
0.2 | 1 | 0.77369 | 0.27382 | 0.74302 | 0.74115 | 2.1143 | 2.1168
Table xx
Kn = | 20 |  | 100:5 |  | Configuration: | P3as
Operating Conditions |  | Error |  | Expanded Uncertainty |  | Coverage Factor
Burden (Ω) | Primary current (% of rated) | Magnitude (µA/A) | Phase (µrad) | Magnitude (µA/A) | Phase (µrad) | Magnitude | Phase
0.2 | 125 | 0.46353 | 0.38472 | 0.093281 | 0.089344 | 1.98 | 2.0851
0.2 | 120 | 0.42024 | 0.39361 | 0.09336 | 0.089366 | 1.98 | 2.0849
0.2 | 100 | 0.4566 | 0.40059 | 0.093582 | 0.089503 | 1.9799 | 2.0838
0.2 | 60 | 0.58385 | 0.42568 | 0.095641 | 0.09012 | 1.9788 | 2.0791
0.2 | 40 | 0.70649 | 0.39401 | 0.098192 | 0.091327 | 1.9779 | 2.0711
0.2 | 20 | 0.82919 | 0.37621 | 0.10736 | 0.09866 | 1.9761 | 2.0398
0.2 | 10 | 0.9109 | 0.36213 | 0.13555 | 0.12691 | 1.9832 | 2.0098
0.2 | 5 | 0.99082 | 0.35427 | 0.21196 | 0.20455 | 2.0115 | 2.0237
0.2 | 1 | 1.0382 | 0.35089 | 0.93488 | 0.93395 | 2.0495 | 2.0499
Table 4
Kn = | 20 |  | 100:5 |  | Configuration: | P3as
Operating Conditions |  | Error |  | Expanded Uncertainty |  | Coverage Factor
Burden (Ω) | Primary current (% of rated) | Magnitude (µA/A) | Phase (µrad) | Magnitude (µA/A) | Phase (µrad) | Magnitude | Phase
0.2 | 125 | 0.33484 | 0.29901 | 0.080083 | 0.084815 | 1.9874 | 2.1286
0.2 | 120 | 0.32179 | 0.30178 | 0.08009 | 0.084841 | 1.9874 | 2.1283
0.2 | 100 | 0.36862 | 0.30313 | 0.080606 | 0.084929 | 1.9869 | 2.1272
0.2 | 60 | 0.47795 | 0.3085 | 0.082545 | 0.085202 | 1.9855 | 2.1236
0.2 | 40 | 0.56816 | 0.28261 | 0.084935 | 0.085902 | 1.9843 | 2.1156
0.2 | 20 | 0.66231 | 0.26777 | 0.092987 | 0.090248 | 1.9816 | 2.0819
0.2 | 10 | 0.7245 | 0.26298 | 0.11438 | 0.11022 | 1.9902 | 2.0343
0.2 | 5 | 0.79182 | 0.26229 | 0.17266 | 0.16735 | 2.035 | 2.0561
0.2 | 1 | 0.84616 | 0.25735 | 0.75979 | 0.76023 | 2.0985 | 2.0981
Table 5
Kn = | 4 |  | 20:5 |  | Configuration: | P1bs
Operating Conditions |  | Error |  | Expanded Uncertainty |  | Coverage Factor
Burden (Ω) | Primary current (% of rated) | Magnitude (µA/A) | Phase (µrad) | Magnitude (µA/A) | Phase (µrad) | Magnitude | Phase
0.2 | 125 | 0.14885 | 0.20949 | 0.061542 | 0.080444 | 2.0069 | 2.2052
0.2 | 120 | 0.1299 | 0.21266 | 0.06159 | 0.080464 | 2.0068 | 2.2047
0.2 | 100 | 0.16778 | 0.2126 | 0.06222 | 0.080522 | 2.0052 | 2.2032
0.2 | 60 | 0.24443 | 0.20919 | 0.064494 | 0.08064 | 2.0021 | 2.1999
0.2 | 40 | 0.30365 | 0.18902 | 0.067072 | 0.08099 | 1.9999 | 2.1915
0.2 | 20 | 0.36689 | 0.17659 | 0.073919 | 0.083773 | 1.9969 | 2.1467
0.2 | 10 | 0.40188 | 0.17316 | 0.093335 | 0.098035 | 2.0078 | 2.0736
0.2 | 5 | 0.45283 | 0.16245 | 0.14208 | 0.14225 | 2.0822 | 2.1035
0.2 | 1 | 0.46937 | 0.15265 | 0.62922 | 0.62932 | 2.2022 | 2.202
Table 6
Kn = | 20 |  | 100:5 |  | Configuration: | P1bp
Operating Conditions |  | Error |  | Expanded Uncertainty |  | Coverage Factor
Burden (Ω) | Primary current (% of rated) | Magnitude (µA/A) | Phase (µrad) | Magnitude (µA/A) | Phase (µrad) | Magnitude | Phase
0.2 | 125 | 0.16854 | 0.21318 | 0.072703 | 0.08264 | 1.9934 | 2.1599
0.2 | 120 | 0.14958 | 0.21635 | 0.072744 | 0.082662 | 1.9934 | 2.1595
0.2 | 100 | 0.18746 | 0.21629 | 0.073282 | 0.082725 | 1.9926 | 2.1584
0.2 | 60 | 0.26411 | 0.21288 | 0.075222 | 0.082857 | 1.9909 | 2.156
0.2 | 40 | 0.32334 | 0.19271 | 0.077434 | 0.083236 | 1.9897 | 2.1499
0.2 | 20 | 0.38657 | 0.18028 | 0.083404 | 0.086139 | 1.988 | 2.1167
0.2 | 10 | 0.42157 | 0.17685 | 0.10085 | 0.1003 | 1.9969 | 2.061
0.2 | 5 | 0.47252 | 0.16614 | 0.14661 | 0.14373 | 2.0639 | 2.0951
0.2 | 1 | 0.48906 | 0.15633 | 0.62999 | 0.62958 | 2.1997 | 2.2011
Table 7
Kn = | 20 |  | 100:5 |  | Configuration: | P2bs
Operating Conditions |  | Error |  | Expanded Uncertainty |  | Coverage Factor
Burden (Ω) | Primary current (% of rated) | Magnitude (µA/A) | Phase (µrad) | Magnitude (µA/A) | Phase (µrad) | Magnitude | Phase
0.2 | 125 | 0.2906 | 0.30719 | 0.086963 | 0.087055 | 1.9831 | 2.1044
0.2 | 120 | 0.2367 | 0.30886 | 0.087072 | 0.087092 | 1.983 | 2.104
0.2 | 100 | 0.26526 | 0.31404 | 0.087356 | 0.087206 | 1.9828 | 2.1029
0.2 | 60 | 0.36637 | 0.32548 | 0.089394 | 0.087682 | 1.9816 | 2.0983
0.2 | 40 | 0.45299 | 0.2981 | 0.091697 | 0.088534 | 1.9807 | 2.091
0.2 | 20 | 0.53887 | 0.2829 | 0.099768 | 0.094389 | 1.979 | 2.0578
0.2 | 10 | 0.5935 | 0.27538 | 0.12512 | 0.11871 | 1.9865 | 2.0199
0.2 | 5 | 0.65572 | 0.26463 | 0.19312 | 0.18676 | 2.0209 | 2.0361
0.2 | 1 | 0.68986 | 0.2529 | 0.84552 | 0.84451 | 2.072 | 2.0728
Table 8
Kn = | 100 |  | 500:5 |  | Configuration: | P3ap
Operating Conditions |  | Error |  | Expanded Uncertainty |  | Coverage Factor
Burden (Ω) | Primary current (% of rated) | Magnitude (µA/A) | Phase (µrad) | Magnitude (µA/A) | Phase (µrad) | Magnitude | Phase
0.2 | 125 | 0.27351 | 0.19172 | 0.10318 | 0.14717 | 2.0375 | 2.2023
0.2 | 120 | 0.26047 | 0.19449 | 0.10319 | 0.14718 | 2.0375 | 2.2021
0.2 | 100 | 0.30729 | 0.19584 | 0.10356 | 0.14722 | 2.0362 | 2.2015
0.2 | 60 | 0.41662 | 0.2012 | 0.10498 | 0.14736 | 2.0319 | 2.1997
0.2 | 40 | 0.50683 | 0.17532 | 0.10675 | 0.14769 | 2.0272 | 2.1951
0.2 | 20 | 0.60098 | 0.16048 | 0.11293 | 0.14976 | 2.0144 | 2.1714
0.2 | 10 | 0.66318 | 0.15569 | 0.13043 | 0.16043 | 2.0022 | 2.1041
0.2 | 5 | 0.73049 | 0.155 | 0.18269 | 0.20004 | 2.0249 | 2.0521
0.2 | 1 | 0.78483 | 0.15006 | 0.76176 | 0.76632 | 2.0964 | 2.0921
Table 9
Kn = | 5 |  | 25:5 |  | Configuration: | P2as
Operating Conditions |  | Error |  | Expanded Uncertainty |  | Coverage Factor
Burden (Ω) | Primary current (% of rated) | Magnitude (µA/A) | Phase (µrad) | Magnitude (µA/A) | Phase (µrad) | Magnitude | Phase
0.2 | 125 | 0.41275 | 0.50892 | 0.14509 | 0.23762 | 2.1743 | 2.3669
0.2 | 120 | 0.36251 | 0.51086 | 0.14514 | 0.23762 | 2.1737 | 2.3668
0.2 | 100 | 0.40287 | 0.51601 | 0.14525 | 0.23763 | 2.1725 | 2.3663
0.2 | 60 | 0.52233 | 0.52392 | 0.14608 | 0.23768 | 2.1639 | 2.3645
0.2 | 40 | 0.61505 | 0.50049 | 0.14699 | 0.23775 | 2.1553 | 2.3614
0.2 | 20 | 0.71263 | 0.48946 | 0.15019 | 0.23835 | 2.13 | 2.3406
0.2 | 10 | 0.7734 | 0.48349 | 0.16205 | 0.2422 | 2.077 | 2.2668
0.2 | 5 | 0.83977 | 0.47748 | 0.20179 | 0.26249 | 2.0413 | 2.1461
0.2 | 1 | 0.87991 | 0.4581 | 0.74896 | 0.75962 | 2.1073 | 2.0981
Table 10
Kn = | 10 |  | 50:5 |  | Configuration: | P2asp
Operating Conditions |  | Error |  | Expanded Uncertainty |  | Coverage Factor
Burden (Ω) | Primary current (% of rated) | Magnitude (µA/A) | Phase (µrad) | Magnitude (µA/A) | Phase (µrad) | Magnitude | Phase
0.2 | 125 | 0.31343 | 0.35079 | 0.17407 | 0.28212 | 2.1101 | 2.1799
0.2 | 120 | 0.26319 | 0.35273 | 0.17412 | 0.28213 | 2.1099 | 2.1798
0.2 | 100 | 0.30355 | 0.35788 | 0.17422 | 0.28215 | 2.1093 | 2.1797
0.2 | 60 | 0.42301 | 0.36579 | 0.175 | 0.28224 | 2.1054 | 2.1791
0.2 | 40 | 0.51573 | 0.34237 | 0.17586 | 0.28239 | 2.1014 | 2.1782
0.2 | 20 | 0.61331 | 0.33133 | 0.17878 | 0.2835 | 2.089 | 2.1717
0.2 | 10 | 0.67408 | 0.32536 | 0.18916 | 0.2887 | 2.0591 | 2.1466
0.2 | 5 | 0.74045 | 0.31936 | 0.22364 | 0.30814 | 2.0311 | 2.0951
0.2 | 1 | 0.78059 | 0.29997 | 0.75392 | 0.77404 | 2.1018 | 2.0855
Table 11
Kn = | 2 |  | 10:5 |  | Configuration: | P1asp
Operating Conditions |  | Error |  | Expanded Uncertainty |  | Coverage Factor
Burden (Ω) | Primary current (% of rated) | Magnitude (µA/A) | Phase (µrad) | Magnitude (µA/A) | Phase (µrad) | Magnitude | Phase
0.2 | 125 | 0.20461 | 0.22213 | 0.051545 | 0.065132 | 2.0275 | 2.1895
0.2 | 120 | 0.20829 | 0.22321 | 0.051618 | 0.065152 | 2.0273 | 2.189
0.2 | 100 | 0.22933 | 0.22533 | 0.052151 | 0.065203 | 2.0253 | 2.1877
0.2 | 60 | 0.31103 | 0.21392 | 0.054766 | 0.065212 | 2.0192 | 2.1867
0.2 | 40 | 0.36989 | 0.20294 | 0.057328 | 0.065362 | 2.0168 | 2.1825
0.2 | 20 | 0.44162 | 0.19656 | 0.062966 | 0.067138 | 2.0139 | 2.151
0.2 | 10 | 0.4883 | 0.18899 | 0.07461 | 0.075609 | 2.0339 | 2.115
0.2 | 5 | 0.53907 | 0.18402 | 0.11107 | 0.10996 | 2.1547 | 2.2025
0.2 | 1 | 0.5739 | 0.19446 | 0.49559 | 0.4958 | 2.4668 | 2.4963
Table 12
Kn = | 40 |  | 200:5 |  | Configuration: | P2bp
Operating Conditions |  | Error |  | Expanded Uncertainty |  | Coverage Factor
Burden (Ω) | Primary current (% of rated) | Magnitude (µA/A) | Phase (µrad) | Magnitude (µA/A) | Phase (µrad) | Magnitude | Phase
0.2 | 125 | 0.29053 | 0.30661 | 0.21408 | 0.21242 | 1.9606 | 1.9628
0.2 | 120 | 0.23663 | 0.30829 | 0.21412 | 0.21244 | 1.9606 | 1.9628
0.2 | 100 | 0.26519 | 0.31346 | 0.21424 | 0.21249 | 1.9606 | 1.9628
0.2 | 60 | 0.3663 | 0.3249 | 0.21508 | 0.21273 | 1.9606 | 1.9628
0.2 | 40 | 0.45292 | 0.29752 | 0.21605 | 0.21315 | 1.9606 | 1.9628
0.2 | 20 | 0.5388 | 0.28232 | 0.21958 | 0.21594 | 1.9607 | 1.9628
0.2 | 10 | 0.59343 | 0.2748 | 0.23188 | 0.22778 | 1.9621 | 1.9638
0.2 | 5 | 0.65565 | 0.26406 | 0.27298 | 0.26806 | 1.9735 | 1.9754
0.2 | 1 | 0.68979 | 0.25232 | 0.86511 | 0.86409 | 2.0591 | 2.0597

## Descriptors
Label | Kn | Ratio
P1as | 1 | 5:5
P1ap | 4 | 20:5
P2ap | 20 | 100:5
P3as | 20 | 100:5
P1bs | 4 | 20:5
P1bp | 20 | 100:5
P2bs | 20 | 100:5
P3ap | 100 | 500:5
P2as | 5 | 25:5
P2asp | 10 | 50:5
P1asp | 2 | 10:5
P2bp | 40 | 200:5