## Data
CALIBRATION OF KDK CT-12C CT s/n 23975
3 | DEMAGNETISE WITH 8.0 V RMS M-L3 if it was not done in the equal ratio spreadsheet or if you suspect that it has been magnetised. | [15 empty] | Formulas changed.
4 | Note that this procedure uses the Zetka + the CH5500. Ramping the CH5500 too quickly or switching the current off or on abruptly can magnetise this CT. (And maybe our 2 stage!!!) | [15 empty] | From the previous
5 | [16 empty] | calibration | [6 empty] | Power coeff of 0.2 ohm shunt is |  |  | 2.11 | ppm/W
6 |  | Ti969 0.1 ohm shunt vs 0.2 paddlewheel. 2 stage CT in cct at 10/5 ratio.
7 | [10 empty] | transformation ratio |  | IVD | Shunt |  | Forward | dv/V | dv/V | alpha.Nt/Ns | Res | Ratio x | Phase | Mean | Mean | Power | Heating | Power | Heating
8 |  | Shunt V | X | Y | SD X | SD Y | SD Shunt | T 0.1ohm | % load | sec cur | Ns | Nt | alpha | Rs | Rt | Reverse | Re | Im |  | ratio | Ns | error | ratio | phase | in 0.2 shunt | ppm | in 0.1 | [9 empty] | Initial setup. IVD capacitor (GR-1412-BC) set to 0
9 | 2018-10-10
10 | 53 Hz | 0.10017 | 1.1518e-07 | 1.4346e-05 | 4.38e-08 | 1.1082e-07 | 1.1191e-07 | 20.44 | 20.0345 | 1.0011 | 2 | 1 | 1.0007 | 0.19999 | 0.10006 | -1 | -1.149e-06 | -0.00014311 | 0.50036 | 0.50036 | 1.0007 | 0.00014321 | 1.0007 | 0.00010878 | 0.050108 | 0.10573 | 0.10022 | -5.6298 | [7 empty] | 53Hz | 0.10017 | 1.1518e-07 | 1.4346e-05 | 4.38e-08 | 1.1082e-07 | 1.1191e-07 | 12 | 2 | 12,2
11 |  | 0.30123 | 2.5059e-06 | 4.5484e-05 | 6.579e-08 | 1.2878e-07 | 6.5353e-08 | 20.44 | 60.2468 | 3.0104 | 2 | 1 | 1.0007 | 0.19999 | 0.10006 | -1 | -8.3127e-06 | -0.00015088 | 0.50036 | 0.50037 | 1.0007 | 0.00015099 | 1.0007 | 0.00011712 | 0.45312 | 0.95609 | 0.90625 | 1.0064 | [7 empty] | 53Hz | 0.30123 | 2.5059e-06 | 4.5484e-05 | 6.579e-08 | 1.2878e-07 | 6.5353e-08 | 13 | 2 | 13,2
12 |  | 0.50087 | 9.8791e-06 | 7.9211e-05 | 1.7811e-07 | 4.4567e-07 | 8.119e-08 | 20.45 | 100.1742 | 5.0055 | 2 | 1 | 1.0007 | 0.19999 | 0.10006 | -1 | -1.971e-05 | -0.00015803 | 0.50036 | 0.50037 | 1.0007 | 0.00015815 | 1.0007 | 0.00012406 | 1.2527 | 2.6433 | 2.5055 | 13.0331 | [7 empty] | 53Hz | 0.50087 | 9.8791e-06 | 7.9211e-05 | 1.7811e-07 | 4.4567e-07 | 8.119e-08 | 14 | 2 | 14,2
13 |  | 0.60004 | 1.6966e-05 | 9.6685e-05 | 2.7014e-07 | 5.4215e-07 | 7.6551e-08 | 20.45 | 120.0077 | 5.9965 | 2 | 1 | 1.0007 | 0.19999 | 0.10006 | -1 | -2.8255e-05 | -0.00016102 | 0.50036 | 0.50038 | 1.0008 | 0.00016113 | 1.0007 | 0.00012716 | 1.7979 | 3.7936 | 3.5958 | 22.8187 | [7 empty] | 53Hz | 0.60004 | 1.6966e-05 | 9.6685e-05 | 2.7014e-07 | 5.4215e-07 | 7.6551e-08 | 14 | 2 | 14,2
14 | Reverse the SR830 A & B inputs | 0.10039 | 5.8854e-06 | -7.463e-06 | 1.7842e-08 | 6.8587e-08 | 5.4644e-08 | 20.47 | 20.0785 | 1.0033 | 2 | 1 | 1.0007 | 0.19999 | 0.10006 | 1 | 5.8582e-05 | -7.4285e-05 | 0.50036 | 0.50033 | 1.0007 | 7.4338e-05 |  | intercept = | -6.479 | [9 empty] | A - B reversed | 53Hz | 0.10039 | 5.8854e-06 | -7.463e-06 | 1.7842e-08 | 6.8587e-08 | 5.4644e-08 | 11 | 2 | 11,2
15 |  | 0.30069 | 1.6304e-05 | -2.5034e-05 | 1.5021e-07 | 4.8408e-08 | 3.2444e-08 | 20.48 | 60.1381 | 3.005 | 2 | 1 | 1.0007 | 0.19999 | 0.10006 | 1 | 5.4182e-05 | -8.3195e-05 | 0.50036 | 0.50033 | 1.0007 | 8.3255e-05 |  | slope = | 8.0397 | ppm/watt | [8 empty] | A - B reversed | 53Hz | 0.30069 | 1.6304e-05 | -2.5034e-05 | 1.5021e-07 | 4.8408e-08 | 3.2444e-08 | 13 | 2 | 13,2
16 |  | 0.49747 | 2.236e-05 | -4.4754e-05 | 5.1479e-07 | 3.5798e-07 | 5.2748e-08 | 20.5 | 99.4942 | 4.9715 | 2 | 1 | 1.0007 | 0.19999 | 0.10006 | 1 | 4.4914e-05 | -8.9898e-05 | 0.50036 | 0.50034 | 1.0007 | 8.9963e-05 |  |  | Mean ratio at 53 Hz: | 1.0007 | [8 empty] | A - B reversed | 53Hz | 0.49747 | 2.236e-05 | -4.4754e-05 | 5.1479e-07 | 3.5798e-07 | 5.2748e-08 | 13 | 2 | 13,2
17 |  | 0.60449 | 2.1899e-05 | -5.6334e-05 | 1.5796e-07 | 4.968e-07 | 5.6139e-08 | 20.52 | 120.8982 | 6.041 | 2 | 1 | 1.0007 | 0.19999 | 0.10006 | 1 | 3.6201e-05 | -9.3125e-05 | 0.50036 | 0.50034 | 1.0007 | 9.3193e-05 |  |  | Mean phase at 53 Hz: | 0.00011928 | [8 empty] | A - B reversed | 53Hz | 0.60449 | 2.1899e-05 | -5.6334e-05 | 1.5796e-07 | 4.968e-07 | 5.6139e-08 | 14 | 2 | 14,2
18
19 | 47 Hz | 0.10071 | -1.5788e-07 | 1.0614e-05 | 5.1866e-08 | 9.0334e-08 | 8.7481e-08 | 20.47 | 20.1413 | 1.0064 | 2 | 1 | 1.0007 | 0.19999 | 0.10006 | -1 | 1.5666e-06 | -0.00010532 | 0.50036 | 0.50036 | 1.0007 | 0.0001054 | 1.0007 | 6.5766e-05 | 0.050644 | 0.10686 | 0.10129 | -8.0358 | [7 empty] | 47Hz | 0.10071 | -1.5788e-07 | 1.0614e-05 | 5.1866e-08 | 9.0334e-08 | 8.7481e-08 | 11 | 2 | 11,2
20 |  | 0.30065 | 1.9494e-06 | 3.4624e-05 | 8.5779e-08 | 1.3153e-07 | 7.2027e-08 | 20.47 | 60.131 | 3.0046 | 2 | 1 | 1.0007 | 0.19999 | 0.10006 | -1 | -6.4793e-06 | -0.00011508 | 0.50036 | 0.50036 | 1.0007 | 0.00011516 | 1.0007 | 7.5504e-05 | 0.45138 | 0.95242 | 0.90277 | -0.51776 | [7 empty] | 47Hz | 0.30065 | 1.9494e-06 | 3.4624e-05 | 8.5779e-08 | 1.3153e-07 | 7.2027e-08 | 13 | 2 | 13,2
21 |  | 0.50137 | 9.2622e-06 | 6.1686e-05 | 3.6751e-06 | 4.041e-07 | 8.77e-08 | 20.47 | 100.2739 | 5.0105 | 2 | 1 | 1.0007 | 0.19999 | 0.10006 | -1 | -1.846e-05 | -0.00012295 | 0.50036 | 0.50037 | 1.0007 | 0.00012303 | 1.0007 | 8.3411e-05 | 1.2552 | 2.6485 | 2.5105 | 12.0903 | [7 empty] | 47Hz | 0.50137 | 9.2622e-06 | 6.1686e-05 | 3.6751e-06 | 4.041e-07 | 8.77e-08 | 14 | 2 | 14,2
22 |  | 0.60236 | 1.6244e-05 | 7.6347e-05 | 7.9236e-07 | 5.5853e-07 | 8.2275e-08 | 20.47 | 120.4714 | 6.0197 | 2 | 1 | 1.0007 | 0.19999 | 0.10006 | -1 | -2.6948e-05 | -0.00012666 | 0.50036 | 0.50037 | 1.0007 | 0.00012675 | 1.0007 | 8.7012e-05 | 1.8118 | 3.823 | 3.6237 | 21.8019 | [7 empty] | 47Hz | 0.60236 | 1.6244e-05 | 7.6347e-05 | 7.9236e-07 | 5.5853e-07 | 8.2275e-08 | 14 | 2 | 14,2
23 | Reverse the SR830 A & B inputs | 0.10099 | 6.132e-06 | -2.639e-06 | 4.3073e-08 | 8.9797e-08 | 6.4521e-08 | 20.47 | 20.1973 | 1.0092 | 2 | 1 | 1.0007 | 0.19999 | 0.10006 | 1 | 6.0677e-05 | -2.6113e-05 | 0.50036 | 0.50033 | 1.0007 | 2.6132e-05 |  | intercept = | -8.581 | [9 empty] | A - B reversed | 47Hz | 0.10099 | 6.132e-06 | -2.639e-06 | 4.3073e-08 | 8.9797e-08 | 6.4521e-08 | 11 | 2 | 11,2
24 |  | 0.30082 | 1.6674e-05 | -1.0784e-05 | 1.6155e-07 | 9.642e-08 | 6.4223e-08 | 20.47 | 60.1637 | 3.0062 | 2 | 1 | 1.0007 | 0.19999 | 0.10006 | 1 | 5.5387e-05 | -3.5822e-05 | 0.50036 | 0.50033 | 1.0007 | 3.5848e-05 |  | slope = | 8.3582 | ppm/watt | [8 empty] | A - B reversed | 47Hz | 0.30082 | 1.6674e-05 | -1.0784e-05 | 1.6155e-07 | 9.642e-08 | 6.4223e-08 | 12 | 2 | 12,2
25 |  | 0.50095 | 2.284e-05 | -2.1935e-05 | 2.5303e-06 | 3.7966e-07 | 7.5887e-08 | 20.48 | 100.1899 | 5.0063 | 2 | 1 | 1.0007 | 0.19999 | 0.10006 | 1 | 4.556e-05 | -4.3755e-05 | 0.50036 | 0.50034 | 1.0007 | 4.3787e-05 |  |  | Mean ratio at 47 Hz: | 1.0007 | [8 empty] | A - B reversed | 47Hz | 0.50095 | 2.284e-05 | -2.1935e-05 | 2.5303e-06 | 3.7966e-07 | 7.5887e-08 | 13 | 2 | 13,2
26 |  | 0.60184 | 2.2275e-05 | -2.8452e-05 | 2.6262e-07 | 5.3038e-07 | 8.7077e-08 | 20.51 | 120.3679 | 6.0145 | 2 | 1 | 1.0007 | 0.19999 | 0.10006 | 1 | 3.6985e-05 | -4.7242e-05 | 0.50036 | 0.50034 | 1.0007 | 4.7276e-05 |  |  | Mean phase at 47 Hz: | 7.7923e-05 | [8 empty] | A - B reversed | 47Hz | 0.60184 | 2.2275e-05 | -2.8452e-05 | 2.6262e-07 | 5.3038e-07 | 8.7077e-08 | 13 | 2 | 13,2
27 | [34 empty] | Repeat A - B reversed | 47Hz | 0.10115 | 6.4216e-06 | -2.714e-06 | 4.3211e-08 | 6.433e-08 | 7.174e-08 | 11 | 2 | 11,2
28 | [23 empty] | Heating in 0.1 ohm = | 8.199 | ppm/watt | [8 empty] | A - B reversed with 1k across 0.1 potential terminals | 47Hz | 0.10112 | 1.6934e-05 | -2.5039e-06 | 1.196e-08 | 6.3074e-08 | 6.1846e-08 | 12 | 2 | 12,2
29 |  | 30/5 range, l - 1200At(5A), vs shunts and 500 AT on 50/5 range. | [24 empty] | mean of 53 and 47 Hz | [7 empty] | Balance with 1 k across. | X005550
30 | [10 empty] | transformation ratio |  | IVD | Shunt |  | Forward | dv/V | dv/V | [6 empty] | heating | heating | heating corrected
31 | 2018-10-10 | [6 empty] | T 0.1ohm | % load | sec cur | Ns | Nt | alpha | Rs | Rt | Reverse | Re | Im |  |  |  |  | mA/A | mrad | 0.1 shunt | 0.2 shunt | mA/A | mrad
32 | 53 Hz | 0.030152 | -1.3056e-07 | 6.9734e-06 | 3.5005e-08 | 1.382e-07 | 1.7797e-07 | 20.67 | 5.0253 | 0.25127 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | -5.1932e-06 | 0.00027737 | 0.99988 | 1.0007 | 0.00011928 | 0.99987 | -0.12704 | 0.15809 | -0.00076984 | -9.6136e-05 | -0.13138 | 0.16524 | [7 empty] | 53Hz | 0.030152 | -1.3056e-07 | 6.9734e-06 | 3.5005e-08 | 1.382e-07 | 1.7797e-07 | 11 | 2
33 | UUT range | 0.060053 | -2.0447e-07 | 1.3502e-05 | 3.3726e-08 | 1.1474e-07 | 1.5084e-07 | 20.67 | 10.0089 | 0.50044 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | -4.0835e-06 | 0.00026964 | 0.99988 | 1.0007 | 0.00011928 | 0.99987 | -0.12593 | 0.15036 | -0.00061609 | -6.768e-05 | -0.13137 | 0.15641 | [7 empty] | 53Hz | 0.060053 | -2.0447e-07 | 1.3502e-05 | 3.3726e-08 | 1.1474e-07 | 1.5084e-07 | 12 | 2
34 | 30/5 | 0.1202 | -5.9607e-07 | 2.6061e-05 | 5.3853e-08 | 2.5184e-07 | 1.9118e-07 | 20.67 | 20.0329 | 1.0016 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | -5.9474e-06 | 0.00026003 | 0.99988 | 1.0007 | 0.00011928 | 0.99987 | -0.12779 | 0.14075 | 1.8941e-06 | 4.6693e-05 | -0.1332 | 0.14615 | [7 empty] | 53Hz | 0.1202 | -5.9607e-07 | 2.6061e-05 | 5.3853e-08 | 2.5184e-07 | 1.9118e-07 | 13 | 2
35 | UUT secondary | 0.23978 | -2.0769e-07 | 4.9604e-05 | 6.2779e-08 | 1.3863e-07 | 1.4922e-07 | 20.67 | 39.9637 | 1.9982 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | -1.0388e-06 | 0.0002481 | 0.99988 | 1.0007 | 0.00011928 | 0.99988 | -0.12289 | 0.12882 | 0.0024558 | 0.00050085 | -0.13045 | 0.1338 | [7 empty] | 53Hz | 0.23978 | -2.0769e-07 | 4.9604e-05 | 6.2779e-08 | 1.3863e-07 | 1.4922e-07 | 13 | 2
36 | l - 1200At(5A) | 0.35946 | 2.5168e-06 | 7.1805e-05 | 6.7085e-08 | 2.6775e-07 | 2.577e-07 | 20.67 | 59.9092 | 2.9955 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | 8.3972e-06 | 0.00023957 | 0.99988 | 1.0007 | 0.00011928 | 0.99989 | -0.11345 | 0.12029 | 0.0065437 | 0.0012574 | -0.12447 | 0.12558 | [7 empty] | 53Hz | 0.35946 | 2.5168e-06 | 7.1805e-05 | 6.7085e-08 | 2.6775e-07 | 2.577e-07 | 14 | 2
37 | Std Range | 0.60059 | 1.4445e-05 | 0.00011564 | 6.7592e-07 | 4.5575e-07 | 2.3269e-07 | 20.67 | 100.0984 | 5.0049 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | 2.8844e-05 | 0.00023092 | 0.99988 | 1.0007 | 0.00011928 | 0.99991 | -0.093003 | 0.11164 | 0.01974 | 0.0036998 | -0.11449 | 0.11781 | [7 empty] | 53Hz | 0.60059 | 1.4445e-05 | 0.00011564 | 6.7592e-07 | 4.5575e-07 | 2.3269e-07 | 15 | 2
38 | 50/5 | 0.7214 | 2.4835e-05 | 0.00013777 | 2.2112e-06 | 4.3534e-07 | 1.8333e-07 | 20.68 | 120.2336 | 6.0117 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | 4.1287e-05 | 0.00022904 | 0.99988 | 1.0007 | 0.00011928 | 0.99992 | -0.080559 | 0.10976 | 0.028844 | 0.0053847 | -0.10945 | 0.1165 | [7 empty] | 53Hz | 0.7214 | 2.4835e-05 | 0.00013777 | 2.2112e-06 | 4.3534e-07 | 1.8333e-07 | 15 | 2
39
40 | 47 Hz | 0.029685 | -3.7688e-07 | 6.1957e-06 | 2.8489e-08 | 9.1634e-08 | 1.1429e-07 | 21.5 | 4.9505 | 0.24752 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | -1.5226e-05 | 0.00025031 | 0.99988 | 1.0007 | 7.7923e-05 | 0.99986 | -0.13707 | 0.17239 | -0.00077143 | -9.6431e-05 | [9 empty] | 47Hz | 0.029685 | -3.7688e-07 | 6.1957e-06 | 2.8489e-08 | 9.1634e-08 | 1.1429e-07 | 11 | 2
41 |  | 0.060225 | -8.0699e-07 | 1.2071e-05 | 3.3846e-07 | 9.127e-08 | 9.1749e-08 | 21.51 | 10.0433 | 0.50217 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | -1.607e-05 | 0.00024037 | 0.99988 | 1.0007 | 7.7923e-05 | 0.99986 | -0.13792 | 0.16245 | -0.00061491 | -6.7462e-05 | [9 empty] | 47Hz | 0.060225 | -8.0699e-07 | 1.2071e-05 | 3.3846e-07 | 9.127e-08 | 9.1749e-08 | 12 | 2
42 |  | 0.11991 | -1.6855e-06 | 2.2943e-05 | 7.3373e-08 | 7.4042e-08 | 9.6448e-08 | 21.52 | 19.9969 | 0.99984 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | -1.6857e-05 | 0.00022947 | 0.99988 | 1.0007 | 7.7923e-05 | 0.99986 | -0.1387 | 0.15154 | -2.0251e-06 | 4.5968e-05 | [9 empty] | 47Hz | 0.11991 | -1.6855e-06 | 2.2943e-05 | 7.3373e-08 | 7.4042e-08 | 9.6448e-08 | 13 | 2
43 |  | 0.24105 | -2.4648e-06 | 4.3557e-05 | 3.493e-08 | 1.2057e-07 | 1.1173e-07 | 21.53 | 40.1992 | 2.01 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | -1.2263e-05 | 0.00021671 | 0.99988 | 1.0007 | 7.7923e-05 | 0.99987 | -0.13411 | 0.13878 | 0.0024907 | 0.0005073 | [9 empty] | 47Hz | 0.24105 | -2.4648e-06 | 4.3557e-05 | 3.493e-08 | 1.2057e-07 | 1.1173e-07 | 13 | 2
44 |  | 0.35965 | -9.1985e-07 | 6.2612e-05 | 6.3655e-06 | 1.5158e-07 | 1.5078e-07 | 21.53 | 59.9773 | 2.9989 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | -3.0673e-06 | 0.00020879 | 0.99988 | 1.0007 | 7.7923e-05 | 0.99988 | -0.12491 | 0.13086 | 0.0065518 | 0.0012589 | [9 empty] | 47Hz | 0.35965 | -9.1985e-07 | 6.2612e-05 | 6.3655e-06 | 1.5158e-07 | 1.5078e-07 | 14 | 2
45 |  | 0.59989 | 8.9814e-06 | 0.00010099 | 2.9676e-06 | 3.9544e-07 | 2.3372e-07 | 21.54 | 100.0403 | 5.002 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | 1.7955e-05 | 0.0002019 | 0.99988 | 1.0007 | 7.7923e-05 | 0.9999 | -0.10389 | 0.12398 | 0.019692 | 0.0036909 | [9 empty] | 47Hz | 0.59989 | 8.9814e-06 | 0.00010099 | 2.9676e-06 | 3.9544e-07 | 2.3372e-07 | 14 | 2
46 |  | 0.72274 | 1.8332e-05 | 0.00012123 | 1.2061e-06 | 5.3789e-07 | 2.0934e-07 | 21.54 | 120.5263 | 6.0263 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | 3.0421e-05 | 0.00020116 | 0.99988 | 1.0007 | 7.7923e-05 | 0.99991 | -0.091426 | 0.12324 | 0.028954 | 0.005405 | [9 empty] | 47Hz | 0.72274 | 1.8332e-05 | 0.00012123 | 1.2061e-06 | 5.3789e-07 | 2.0934e-07 | 15 | 2
47 | 47 Hz | 0.24111 | -1.3363e-07 | 4.3268e-05 | 1.9709e-07 | 1.4973e-07 | 1.2005e-07 | 21.54 | 40.2091 | 2.0105 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | -6.6466e-07 | 0.00021522 | 0.99988 | 1.0007 | 7.7923e-05 | 0.99988 | -0.12251 | 0.13729 | Burden: | R Re | R Im | [8 empty] | 47Hz | 0.24111 | -1.3363e-07 | 4.3268e-05 | 1.9709e-07 | 1.4973e-07 | 1.2005e-07 | 13 | 2
48 | 100 ohms burden test (47Hz) | 0.24133 | -0.00026025 | 3.3018e-05 | 1.2367e-07 | 1.3771e-07 | 1.5923e-07 | 21.54 | 40.2455 | 2.0123 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | -0.0012933 | 0.00016408 | 0.99988 | 1.0007 | 7.7923e-05 | 0.99858 | -1.4152 | 0.086161 |  | 0.12927 | 0.0051132 | [8 empty] | 47Hz | 0.24133 | -0.00026025 | 3.3018e-05 | 1.2367e-07 | 1.3771e-07 | 1.5923e-07 | 16 | 2
49 | 1uF across CT output | 0.24143 | -3.8948e-07 | 3.6029e-05 | 3.1634e-08 | 1.3243e-07 | 1.3359e-07 | 21.54 | 40.2619 | 2.0131 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | -1.9347e-06 | 0.00017897 | 0.99988 | 1.0007 | 7.7923e-05 | 0.99988 | -0.12378 | 0.10105 |  | 0.00012701 | 0.0036245 | [8 empty] | 47Hz | 0.24143 | -3.8948e-07 | 3.6029e-05 | 3.1634e-08 | 1.3243e-07 | 1.3359e-07 | 13 | 2
51 | SR830 reads across 0.1 as above. (The 0.5 is in series on the high side) | [25 empty] | mean of 53 and 47 Hz
52 |  | 30/5 range, l - 1200At(5A), vs shunts and 500 AT on 50/5 range. | [6 empty] | add 0.5 ohms to burden (2 Tinsley 1659 1ohm in parallel) | [15 empty] | heating | heating | heating corrected |  |  |  | Burden error from | 0.6 | ohms | For burden uncert
53 | 2018-10-10 | [21 empty] | mA/A | mrad | 0.1 shunt | 0.2 shunt | mA/A | mrad |  |  | Ratio | Phase
54 | 53 Hz | 0.029914 | -2.4999e-06 | 1.0906e-05 | 2.6477e-08 | 8.2952e-08 | 1.0757e-07 | 20.94 | 4.9885 | 0.24943 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | -0.00010023 | 0.00043724 | 0.99988 | 1.0007 | 0.00011928 | 0.99978 | -0.22207 | 0.31796 | -0.00077066 | -9.6287e-05 | -0.22973 | 0.3363 |  |  | -0.093966 | 0.15274 |  |  |  | 53Hz | 0.029914 | -2.4999e-06 | 1.0906e-05 | 2.6477e-08 | 8.2952e-08 | 1.0757e-07 | 12 | 2
55 | UUT range | 0.059988 | -5.1322e-06 | 2.0718e-05 | 3.2556e-08 | 1.1544e-07 | 1.2132e-07 | 20.94 | 10.0038 | 0.50019 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | -0.0001026 | 0.0004142 | 0.99988 | 1.0007 | 0.00011928 | 0.99978 | -0.22445 | 0.29492 | -0.00061653 | -6.7763e-05 | -0.23134 | 0.30983 |  |  | -0.094126 | 0.13811 |  |  |  | 53Hz | 0.059988 | -5.1322e-06 | 2.0718e-05 | 3.2556e-08 | 1.1544e-07 | 1.2132e-07 | 12 | 2
56 | 30/5 | 0.12003 | -1.061e-05 | 3.8718e-05 | 1.5295e-08 | 1.4091e-07 | 9.3233e-08 | 20.94 | 20.0172 | 1.0009 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | -0.00010601 | 0.00038684 | 0.99988 | 1.0007 | 0.00011928 | 0.99977 | -0.22786 | 0.26757 | -3.551e-07 | 4.6277e-05 | -0.23525 | 0.27999 |  |  | -0.095601 | 0.12116 |  |  |  | 53Hz | 0.12003 | -1.061e-05 | 3.8718e-05 | 1.5295e-08 | 1.4091e-07 | 9.3233e-08 | 13 | 2
57 | UUT secondary | 0.2404 | -2.1058e-05 | 7.1248e-05 | 3.1803e-07 | 1.8518e-07 | 1.8005e-07 | 20.94 | 40.0896 | 2.0045 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | -0.00010506 | 0.00035544 | 0.99988 | 1.0007 | 0.00011928 | 0.99977 | -0.2269 | 0.23616 | 0.0024726 | 0.00050397 | -0.23574 | 0.24682 |  |  | -0.099379 | 0.10255 |  |  |  | 53Hz | 0.2404 | -2.1058e-05 | 7.1248e-05 | 3.1803e-07 | 1.8518e-07 | 1.8005e-07 | 14 | 2
58 | l - 1200At(5A) | 0.36011 | -2.9245e-05 | 0.00010048 | 8.1663e-07 | 9.6633e-08 | 1.5007e-07 | 20.93 | 60.0529 | 3.0026 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | -9.7397e-05 | 0.00033464 | 0.99988 | 1.0007 | 0.00011928 | 0.99978 | -0.21924 | 0.21536 | 0.0065704 | 0.0012624 | -0.23135 | 0.22557 |  |  | -0.10108 | 0.090825 |  |  |  | 53Hz | 0.36011 | -2.9245e-05 | 0.00010048 | 8.1663e-07 | 9.6633e-08 | 1.5007e-07 | 14 | 2
59 | Std Range | 0.60012 | -3.9762e-05 | 0.0001573 | 1.1994e-06 | 4.894e-07 | 6.4515e-07 | 20.93 | 100.0788 | 5.0039 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | -7.9462e-05 | 0.00031435 | 0.99988 | 1.0007 | 0.00011928 | 0.9998 | -0.20131 | 0.19507 | 0.019708 | 0.0036938 | -0.22331 | 0.20354 |  |  | -0.10348 | 0.079704 |  |  |  | 53Hz | 0.60012 | -3.9762e-05 | 0.0001573 | 1.1994e-06 | 4.894e-07 | 6.4515e-07 | 15 | 2
60 | 50/5 | 0.72025 | -3.9924e-05 | 0.00018327 | 5.5637e-06 | 5.3676e-07 | 6.0897e-07 | 20.93 | 120.1116 | 6.0056 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | -6.6478e-05 | 0.00030517 | 0.99988 | 1.0007 | 0.00011928 | 0.99981 | -0.18832 | 0.18589 | 0.02875 | 0.0053672 | -0.21787 | 0.19303 |  |  | -0.10296 | 0.072729 |  |  |  | 53Hz | 0.72025 | -3.9924e-05 | 0.00018327 | 5.5637e-06 | 5.3676e-07 | 6.0897e-07 | 15 | 2
61
62 | 47 Hz | 0.02989 | -2.9135e-06 | 1.0781e-05 | 4.3174e-08 | 2.0258e-07 | 2.185e-07 | 20.95 | 4.9846 | 0.24923 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | -0.0001169 | 0.00043256 | 0.99988 | 1.0007 | 7.7923e-05 | 0.99976 | -0.23875 | 0.35463 | -0.00077074 | -9.6302e-05 | [9 empty] | 47Hz | 0.02989 | -2.9135e-06 | 1.0781e-05 | 4.3174e-08 | 2.0258e-07 | 2.185e-07 | 12 | 2
63 |  | 0.060036 | -5.8812e-06 | 2.0157e-05 | 1.8283e-08 | 1.3948e-07 | 1.2986e-07 | 20.96 | 10.0119 | 0.50059 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | -0.00011748 | 0.00040266 | 0.99988 | 1.0007 | 7.7923e-05 | 0.99976 | -0.23933 | 0.32473 | -0.0006162 | -6.7702e-05 | [9 empty] | 47Hz | 0.060036 | -5.8812e-06 | 2.0157e-05 | 1.8283e-08 | 1.3948e-07 | 1.2986e-07 | 12 | 2
64 |  | 0.12057 | -1.2153e-05 | 3.7232e-05 | 3.0877e-08 | 2.4255e-07 | 1.4497e-07 | 20.99 | 20.1067 | 1.0053 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | -0.00012089 | 0.00037035 | 0.99988 | 1.0007 | 7.7923e-05 | 0.99976 | -0.24274 | 0.29242 | 7.0023e-06 | 4.7638e-05 | [9 empty] | 47Hz | 0.12057 | -1.2153e-05 | 3.7232e-05 | 3.0877e-08 | 2.4255e-07 | 1.4497e-07 | 13 | 2
65 |  | 0.24032 | -2.3803e-05 | 6.7207e-05 | 9.9923e-09 | 1.4854e-07 | 1.7253e-07 | 21.02 | 40.0767 | 2.0038 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | -0.00011879 | 0.00033539 | 0.99988 | 1.0007 | 7.7923e-05 | 0.99976 | -0.24063 | 0.25747 | 0.0024705 | 0.00050357 | [9 empty] | 47Hz | 0.24032 | -2.3803e-05 | 6.7207e-05 | 9.9923e-09 | 1.4854e-07 | 1.7253e-07 | 14 | 2
66 |  | 0.36068 | -3.3383e-05 | 9.4348e-05 | 9.4062e-08 | 2.0168e-07 | 1.9277e-07 | 21.04 | 60.1492 | 3.0075 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | -0.000111 | 0.00031371 | 0.99988 | 1.0007 | 7.7923e-05 | 0.99977 | -0.23285 | 0.23579 | 0.0065942 | 0.0012668 | [9 empty] | 47Hz | 0.36068 | -3.3383e-05 | 9.4348e-05 | 9.4062e-08 | 2.0168e-07 | 1.9277e-07 | 14 | 2
67 |  | 0.59995 | -4.5738e-05 | 0.00014503 | 1.8084e-06 | 4.9633e-07 | 7.4821e-07 | 21.05 | 100.0492 | 5.0025 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | -9.1431e-05 | 0.00028993 | 0.99988 | 1.0007 | 7.7923e-05 | 0.99979 | -0.21328 | 0.212 | 0.019696 | 0.0036916 | [9 empty] | 47Hz | 0.59995 | -4.5738e-05 | 0.00014503 | 1.8084e-06 | 4.9633e-07 | 7.4821e-07 | 15 | 2
68 |  | 0.72111 | -4.7385e-05 | 0.00016721 | 2.1419e-06 | 6.9621e-07 | 9.5271e-07 | 21.07 | 120.255 | 6.0128 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | -7.8808e-05 | 0.00027809 | 0.99988 | 1.0007 | 7.7923e-05 | 0.9998 | -0.20065 | 0.20017 | 0.02882 | 0.0053803 | [9 empty] | 47Hz | 0.72111 | -4.7385e-05 | 0.00016721 | 2.1419e-06 | 6.9621e-07 | 9.5271e-07 | 15 | 2
69 | 47 Hz | 0.06048 | -5.4659e-06 | 2.0101e-05 | 6.9555e-08 | 1.0833e-07 | 9.7405e-08 | 21.09 | 10.0859 | 0.50429 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | -0.00010839 | 0.00039859 | 0.99988 | 1.0007 | 7.7923e-05 | 0.99977 | -0.23023 | 0.32067 | Burden: | R Re | R Im | [8 empty] | 47Hz | 0.06048 | -5.4659e-06 | 2.0101e-05 | 6.9555e-08 | 1.0833e-07 | 9.7405e-08 | 12 | 2
70 | 100 ohms burden test (47Hz) | 0.060642 | -0.00031998 | -2.261e-05 | 3.9671e-08 | 1.4598e-07 | 1.3005e-07 | 21.11 | 10.1129 | 0.50564 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | -0.0063281 | -0.00044715 | 0.99988 | 1.0007 | 7.7923e-05 | 0.99355 | -6.45 | -0.52507 |  | 0.62198 | 0.084574 | [8 empty] | 47Hz | 0.060642 | -0.00031998 | -2.261e-05 | 3.9671e-08 | 1.4598e-07 | 1.3005e-07 | 16 | 2
| 1uF across CT output | 0.060569 | -4.4495e-06 | 1.0974e-05 | 5.4485e-08 | 1.224e-07 | 8.9825e-08 | 21.13 | 10.1007 | 0.50504 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | -8.8102e-05 | 0.0002173 | 0.99988 | 1.0007 | 7.7923e-05 | 0.99979 | -0.20995 | 0.13938 |  | 0.61391 | 0.068692 | [8 empty] | 47Hz | 0.060569 | -4.4495e-06 | 1.0974e-05 | 5.4485e-08 | 1.224e-07 | 8.9825e-08 | 12 | 2
72
73 | SR830 reads across 0.1 as above. (The 0.05 is in series on the high side) | [25 empty] | mean of 53 and 47 Hz |  |  | Clamp
74 |  | 30/5 range, l - 1200At(5A), vs shunts and 500 AT on 50/5 range. | [6 empty] | add 0.05 ohms to burden | [15 empty] | heating | heating | heating corrected |  |  | meter
75 | 2018-10-10 | [21 empty] | mA/A | mrad | 0.1 shunt | 0.2 shunt | mA/A | mrad |  | reading
76 | 53 Hz | 0.0299 | -2.006e-07 | 7.0439e-06 | 3.4943e-08 | 1.3524e-07 | 1.3519e-07 | 21.01 | 4.9862 | 0.24931 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | -8.0463e-06 | 0.00028253 | 0.99988 | 1.0007 | 0.00011928 | 0.99987 | -0.12989 | 0.16326 | -0.0007707 | -9.6296e-05 | -0.13279 | 0.17552 |  | 1.495 | [5 empty] | 53Hz | 0.0299 | -2.006e-07 | 7.0439e-06 | 3.4943e-08 | 1.3524e-07 | 1.3519e-07 | 11 | 2
77 | UUT range | 0.060313 | -5.2538e-07 | 1.3822e-05 | 3.3118e-08 | 2.8027e-07 | 2.1096e-07 | 21 | 10.058 | 0.5029 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | -1.0447e-05 | 0.00027484 | 0.99988 | 1.0007 | 0.00011928 | 0.99987 | -0.13229 | 0.15556 | -0.00061431 | -6.7351e-05 | -0.13568 | 0.16549 |  | 3.016 | [5 empty] | 53Hz | 0.060313 | -5.2538e-07 | 1.3822e-05 | 3.3118e-08 | 2.8027e-07 | 2.1096e-07 | 12 | 2
78 | 30/5 | 0.12018 | -1.1151e-06 | 2.6457e-05 | 2.5344e-08 | 3.0656e-07 | 5.6504e-07 | 21 | 20.0409 | 1.002 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | -1.1128e-05 | 0.00026403 | 0.99988 | 1.0007 | 0.00011928 | 0.99987 | -0.13297 | 0.14475 | 1.5896e-06 | 4.6637e-05 | -0.13736 | 0.15395 |  | 6.01 | [5 empty] | 53Hz | 0.12018 | -1.1151e-06 | 2.6457e-05 | 2.5344e-08 | 3.0656e-07 | 5.6504e-07 | 13 | 2
79 | UUT secondary | 0.24005 | -1.4214e-06 | 5.0256e-05 | 6.8918e-08 | 2.822e-07 | 2.0006e-07 | 21 | 40.0321 | 2.0016 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | -7.1012e-06 | 0.00025108 | 0.99988 | 1.0007 | 0.00011928 | 0.99987 | -0.12895 | 0.1318 | 0.0024632 | 0.00050222 | -0.13533 | 0.14007 |  | 12.01 | [5 empty] | 53Hz | 0.24005 | -1.4214e-06 | 5.0256e-05 | 6.8918e-08 | 2.822e-07 | 2.0006e-07 | 13 | 2
80 | l - 1200At(5A) | 0.36021 | 3.9943e-07 | 7.2839e-05 | 2.0531e-07 | 1.17e-07 | 1.0937e-07 | 21 | 60.0701 | 3.0035 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | 1.3299e-06 | 0.00024251 | 0.99988 | 1.0007 | 0.00011928 | 0.99988 | -0.12052 | 0.12324 | 0.0065747 | 0.0012631 | -0.13002 | 0.13112 |  | 18.03 | [5 empty] | 53Hz | 0.36021 | 3.9943e-07 | 7.2839e-05 | 2.0531e-07 | 1.17e-07 | 1.0937e-07 | 14 | 2
81 | Std Range | 0.60049 | 1.0137e-05 | 0.00011794 | 7.2373e-06 | 3.9202e-07 | 1.538e-07 | 20.99 | 100.1397 | 5.007 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | 2.0245e-05 | 0.00023555 | 0.99988 | 1.0007 | 0.00011928 | 0.9999 | -0.1016 | 0.11627 | 0.019733 | 0.0036984 | -0.12156 | 0.1239 |  | 30.06 | [5 empty] | 53Hz | 0.60049 | 1.0137e-05 | 0.00011794 | 7.2373e-06 | 3.9202e-07 | 1.538e-07 | 15 | 2
82 | 50/5 | 0.72195 | 1.9248e-05 | 0.00014122 | 7.0453e-07 | 5.5652e-07 | 2.5832e-07 | 20.99 | 120.3952 | 6.0198 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | 3.1975e-05 | 0.0002346 | 0.99988 | 1.0007 | 0.00011928 | 0.99991 | -0.089871 | 0.11532 | 0.028889 | 0.005393 | -0.11701 | 0.12312 |  | 36.1 | [5 empty] | 53Hz | 0.72195 | 1.9248e-05 | 0.00014122 | 7.0453e-07 | 5.5652e-07 | 2.5832e-07 | 15 | 2
83
84 | 47 Hz | 0.029894 | -3.7886e-07 | 6.6231e-06 | 3.4931e-08 | 1.3525e-07 | 1.7857e-07 | 20.99 | 4.9853 | 0.24927 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | -1.5199e-05 | 0.00026571 | 0.99988 | 1.0007 | 7.7923e-05 | 0.99986 | -0.13705 | 0.18778 | -0.00077072 | -9.6299e-05 |  |  |  | 1.494 | [5 empty] | 47Hz | 0.029894 | -3.7886e-07 | 6.6231e-06 | 3.4931e-08 | 1.3525e-07 | 1.7857e-07 | 11 | 2
|  | 0.060299 | -9.2068e-07 | 1.2737e-05 | 4.0956e-08 | 1.3717e-07 | 1.0466e-07 | 21.01 | 10.0557 | 0.50278 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | -1.8312e-05 | 0.00025334 | 0.99988 | 1.0007 | 7.7923e-05 | 0.99986 | -0.14016 | 0.17541 | -0.0006144 | -6.7368e-05 |  |  |  | 3.014 | [5 empty] | 47Hz | 0.060299 | -9.2068e-07 | 1.2737e-05 | 4.0956e-08 | 1.3717e-07 | 1.0466e-07 | 12 | 2
|  | 0.12047 | -2.0077e-06 | 2.4216e-05 | 5.3983e-08 | 9.3407e-08 | 8.801e-08 | 21.03 | 20.0901 | 1.0045 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | -1.9987e-05 | 0.00024108 | 0.99988 | 1.0007 | 7.7923e-05 | 0.99986 | -0.14183 | 0.16315 | 5.6388e-06 | 4.7386e-05 |  |  |  | 6.03 | [5 empty] | 47Hz | 0.12047 | -2.0077e-06 | 2.4216e-05 | 5.3983e-08 | 9.3407e-08 | 8.801e-08 | 13 | 2
|  | 0.23986 | -3.1899e-06 | 4.5254e-05 | 4.5043e-08 | 1.269e-07 | 1.0057e-07 | 21.05 | 40.0004 | 2 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | -1.5949e-05 | 0.00022627 | 0.99988 | 1.0007 | 7.7923e-05 | 0.99986 | -0.1378 | 0.14834 | 0.002458 | 0.00050126 |  |  |  | 11.99 | [5 empty] | 47Hz | 0.23986 | -3.1899e-06 | 4.5254e-05 | 4.5043e-08 | 1.269e-07 | 1.0057e-07 | 13 | 2
|  | 0.36003 | -2.1166e-06 | 6.5121e-05 | 1.0369e-07 | 1.5323e-07 | 1.3043e-07 | 21.07 | 60.0406 | 3.002 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | -7.0505e-06 | 0.00021692 | 0.99988 | 1.0007 | 7.7923e-05 | 0.99987 | -0.1289 | 0.139 | 0.0065674 | 0.0012618 |  |  |  | 18.01 | [5 empty] | 47Hz | 0.36003 | -2.1166e-06 | 6.5121e-05 | 1.0369e-07 | 1.5323e-07 | 1.3043e-07 | 14 | 2
|  | 0.60083 | 6.2129e-06 | 0.00010494 | 7.9643e-07 | 3.5938e-07 | 2.2718e-07 | 21.08 | 100.1974 | 5.0099 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | 1.2401e-05 | 0.00020946 | 0.99988 | 1.0007 | 7.7923e-05 | 0.99989 | -0.10945 | 0.13154 | 0.019757 | 0.0037028 |  |  |  | 30.01 | [5 empty] | 47Hz | 0.60083 | 6.2129e-06 | 0.00010494 | 7.9643e-07 | 3.5938e-07 | 2.2718e-07 | 14 | 2
|  | 0.72194 | 1.4867e-05 | 0.00012572 | 3.964e-06 | 4.7896e-07 | 1.4257e-07 | 21.09 | 120.3938 | 6.0197 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | 2.4697e-05 | 0.00020884 | 0.99988 | 1.0007 | 7.7923e-05 | 0.9999 | -0.097149 | 0.13092 | 0.028889 | 0.0053929 |  |  |  | 36.1 | Observed while adding resistor or capacitor |  |  |  |  | 47Hz | 0.72194 | 1.4867e-05 | 0.00012572 | 3.964e-06 | 4.7896e-07 | 1.4257e-07 | 15 | 2
| 47 Hz | 0.2416 | -9.6579e-07 | 4.5265e-05 | 2.3757e-07 | 1.5444e-07 | 1.5675e-07 | 21.12 | 40.2905 | 2.0145 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | -4.7941e-06 | 0.00022469 | 0.99988 | 1.0007 | 7.7923e-05 | 0.99987 | -0.12664 | 0.14677 | Burden: | R Re | R Im |  |  | Primary Current |  |  |  | 12.09 |  | 47Hz | 0.2416 | -9.6579e-07 | 4.5265e-05 | 2.3757e-07 | 1.5444e-07 | 1.5675e-07 | 13 | 2
| 100 ohms burden test (47Hz) | 0.2419 | -0.00036647 | 1.7941e-05 | 1.44e-07 | 1.2716e-07 | 1.1709e-07 | 21.14 | 40.3397 | 2.017 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | -0.0018169 | 8.8947e-05 | 0.99988 | 1.0007 | 7.7923e-05 | 0.99806 | -1.9388 | 0.011024 |  | 0.18121 | 0.013575 |  | 100 ohms across the output terminals of the CT under test. |  |  |  |  | 12.1 |  | 47Hz | 0.2419 | -0.00036647 | 1.7941e-05 | 1.44e-07 | 1.2716e-07 | 1.1709e-07 | 16 | 2
| 1uF across CT output | 0.24203 | -8.5552e-07 | 3.4934e-05 | 6.6937e-08 | 9.4435e-08 | 1.125e-07 | 21.17 | 40.3626 | 2.0181 | 10 | 6 | 0.83382 | 0.2 | 0.1 | 1 | -4.2392e-06 | 0.0001731 | 0.99988 | 1.0007 | 7.7923e-05 | 0.99987 | -0.12609 | 0.095175 |  | 0.17472 | 0.0018793 |  | 1uF across the output terminals of the CT under test. |  |  |  |  | 12.1 |  | 47Hz | 0.24203 | -8.5552e-07 | 3.4934e-05 | 6.6937e-08 | 9.4435e-08 | 1.125e-07 | 13 | 2
95 | SR830 reads across 0.1 as above. (The 0.5 is in series on the high side) | [25 empty] | mean of 53 and 47 Hz
96 |  | 30/5 range, l - 1200At(5A), vs shunts and 500 AT on 50/5 range. | [6 empty] | add 0.5 ohms to burden (2 Tinsley 1659 1ohm in parallel) | [17 empty] | heating corrected
97 | 2018-09-26 | [25 empty] | mA/A | mrad
98 | 53 Hz | 0.024717 | -1.0227e-07 | 1.3079e-05 | 1.7768e-08 | 9.6213e-08 | 1.3284e-07 | 20.73 | 4.9433 | 0.24717 | 10 | 6 | 0.83385 | 0.2 | 0.1 | 1 | -4.9621e-06 | 0.00063458 | 0.99991 | 1.0007 | 0.00011928 | 0.99991 | -0.092033 | 0.5153 | -0.00078684 | -9.9282e-05 | -0.10347 | 0.54294
99 | UUT range | 0.050558 | -7.4184e-07 | 2.526e-05 | 1.1003e-08 | 2.6832e-07 | 3.3894e-08 | 20.73 | 10.1117 | 0.50558 | 10 | 6 | 0.83385 | 0.2 | 0.1 | 1 | -1.7597e-05 | 0.00059917 | 0.99991 | 1.0007 | 0.00011928 | 0.9999 | -0.10467 | 0.4799 | -0.00067595 | -7.876e-05 | -0.1098 | 0.50974
100 | 30/5 | 0.10016 | -2.1572e-06 | 4.7776e-05 | 2.1723e-08 | 8.625e-08 | 3.0282e-07 | 20.73 | 20.0324 | 1.0016 | 10 | 6 | 0.83385 | 0.2 | 0.1 | 1 | -2.5829e-05 | 0.00057203 | 0.99991 | 1.0007 | 0.00011928 | 0.99989 | -0.1129 | 0.45275 | -0.00024974 | 1.1466e-07 | -0.12595 | 0.46718
101 | UUT secondary | 0.20007 | -6.2286e-06 | 8.9896e-05 | 5.6152e-08 | 2.3651e-08 | 3.411e-08 | 20.73 | 40.0138 | 2.0007 | 10 | 6 | 0.83385 | 0.2 | 0.1 | 1 | -3.7336e-05 | 0.00053886 | 0.99991 | 1.0007 | 0.00011928 | 0.99988 | -0.12441 | 0.41958 | 0.0014602 | 0.00031656 | -0.13914 | 0.43562
102 | l - 1200At(5A) | 0.3011 | -9.3119e-06 | 0.00013942 | 2.8373e-06 | 4.5163e-07 | 1.1773e-06 | 20.73 | 60.2202 | 3.011 | 10 | 6 | 0.83385 | 0.2 | 0.1 | 1 | -3.7089e-05 | 0.0005553 | 0.99991 | 1.0007 | 0.00011928 | 0.99988 | -0.12416 | 0.43602 | 0.0043468 | 0.00085075 | -0.14155 | 0.43406
103 | Std Range | 0.50238 | -1.049e-05 | 0.00034185 | 2.7175e-06 | 2.9651e-09 | 2.8499e-05 | 20.73 | 100.4756 | 5.0238 | 10 | 6 | 0.83385 | 0.2 | 0.1 | 1 | -2.5042e-05 | 0.00081604 | 0.99991 | 1.0007 | 0.00011928 | 0.99989 | -0.11211 | 0.69676 | 0.013566 | 0.0025569 | -0.14065 | 0.80546
104 | 50/5 | 0.60116 | -6.3295e-06 | 0.00055334 | 5.2795e-06 | 2.0056e-06 | 2.4053e-05 | 20.73 | 120.2329 | 6.0116 | 10 | 6 | 0.83385 | 0.2 | 0.1 | 1 | -1.2627e-05 | 0.0011039 | 0.99991 | 1.0007 | 0.00011928 | 0.9999 | -0.099698 | 0.98458 | 0.019781 | 0.003707 | -0.12729 | 0.94012
105
106 | 47 Hz | 0.024976 | -6.0847e-07 | 1.3506e-05 | 1.8099e-08 | 9.7463e-08 | 6.7488e-08 | 20.66 | 4.9952 | 0.24976 | 10 | 6 | 0.83385 | 0.2 | 0.1 | 1 | -2.9217e-05 | 0.0006485 | 0.99991 | 1.0007 | 7.7923e-05 | 0.99988 | -0.11629 | 0.57057
107 |  | 0.050348 | -1.22e-06 | 2.5925e-05 | 4.3932e-09 | 1.5138e-07 | 2.5969e-07 | 20.66 | 10.0697 | 0.50348 | 10 | 6 | 0.83385 | 0.2 | 0.1 | 1 | -2.906e-05 | 0.00061752 | 0.99991 | 1.0007 | 7.7923e-05 | 0.99988 | -0.11613 | 0.53959
108 |  | 0.10064 | -4.4002e-06 | 4.6955e-05 | 1.5166e-08 | 3.4088e-08 | 1.0599e-07 | 20.66 | 20.128 | 1.0064 | 10 | 6 | 0.83385 | 0.2 | 0.1 | 1 | -5.2434e-05 | 0.00055953 | 0.99991 | 1.0007 | 7.7923e-05 | 0.99986 | -0.13951 | 0.48161
109 |  | 0.20029 | -1.0774e-05 | 8.8446e-05 | 6.223e-08 | 1.8978e-07 | 1.0303e-07 | 20.66 | 40.0571 | 2.0029 | 10 | 6 | 0.83385 | 0.2 | 0.1 | 1 | -6.4509e-05 | 0.00052959 | 0.99991 | 1.0007 | 7.7923e-05 | 0.99985 | -0.15158 | 0.45167
110 |  | 0.30014 | -1.6238e-05 | 0.00012764 | 6.0592e-07 | 2.2298e-07 | 3.4985e-08 | 20.66 | 60.0278 | 3.0014 | 10 | 6 | 0.83385 | 0.2 | 0.1 | 1 | -6.4883e-05 | 0.00051002 | 0.99991 | 1.0007 | 7.7923e-05 | 0.99985 | -0.15195 | 0.43209
111 |  | 0.50084 | -2.5095e-05 | 0.00041432 | 4.4474e-06 | 5.233e-07 | 2.4987e-05 | 20.67 | 100.1671 | 5.0084 | 10 | 6 | 0.83385 | 0.2 | 0.1 | 1 | -6.0091e-05 | 0.00099209 | 0.99991 | 1.0007 | 7.7923e-05 | 0.99985 | -0.14716 | 0.91417
112 |  | 0.6021 | -1.79e-05 | 0.0004888 | 3.6658e-06 | 5.1949e-07 | 3.0725e-05 | 20.67 | 120.4191 | 6.021 | 10 | 6 | 0.83385 | 0.2 | 0.1 | 1 | -3.5654e-05 | 0.00097359 | 0.99991 | 1.0007 | 7.7923e-05 | 0.99988 | -0.12272 | 0.89566
113 | 53 Hz | 0.30244 | -7.9442e-06 | 0.00016054 | 2.5587e-06 | 5.1743e-07 | 9.0573e-07 | 20.67 | 60.4879 | 3.0244 | 10 | 6 | 0.83385 | 0.2 | 0.1 | 1 | -3.1501e-05 | 0.00063658 | 0.99991 | 1.0007 | 7.7923e-05 | 0.99988 | -0.11857 | 0.55865 | Burden: | R Re | R Im
114 | 100 ohm burden test, 53 Hz | 0.30513 | -0.0015887 | 0.00035859 | 1.6912e-07 | 1.5692e-06 | 7.7089e-06 | 20.67 | 61.0262 | 3.0513 | 10 | 6 | 0.83385 | 0.2 | 0.1 | 1 | -0.0062441 | 0.0014094 | 0.99991 | 1.0007 | 7.7923e-05 | 0.99367 | -6.3312 | 1.3314 |  | 0.62126 | -0.077277
|  | 0.3034 | -1.5522e-05 | 9.6596e-05 | 5.4539e-07 | 7.887e-08 | 6.6307e-06 | 20.68 | 60.6794 | 3.034 | 10 | 6 | 0.83385 | 0.2 | 0.1 | 1 | -6.1356e-05 | 0.00038182 | 0.99991 | 1.0007 | 7.7923e-05 | 0.99985 | -0.14843 | 0.3039
117 |  | 75/5 range, l - 1200At(5A), vs shunts and 500 AT on 100/5 range. | [22 empty] | heating | heating | heating corrected
| 2018-09-26 | [21 empty] | mA/A | mrad | 0.1 shunt | 0.2 shunt | mA/A | mrad
| 53 Hz | 0.024792 | 1.287e-07 | 4.3346e-06 | 1.639e-07 | 1.2622e-07 | 8.8434e-08 | 20.69 | 4.9584 | 0.24792 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | 7.782e-06 | 0.0002621 | 0.9999 | 1.0007 | 0.00011928 | 0.99991 | -0.093079 | 0.14282 | -0.00079924 | -9.9243e-05 | -0.095882 | 0.15972
| UUT range | 0.050305 | 5.6257e-08 | 8.4152e-06 | 2.6427e-07 | 8.9036e-08 | 1.2274e-07 | 20.69 | 10.0609 | 0.50305 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | 1.6765e-06 | 0.00025078 | 0.9999 | 1.0007 | 0.00011928 | 0.9999 | -0.099185 | 0.1315 | -0.00072934 | -7.903e-05 | -0.099721 | 0.15126
| 75/5 | 0.10031 | 2.1104e-08 | 1.6489e-05 | 1.8888e-08 | 4.1979e-08 | 1.661e-07 | 20.69 | 20.0617 | 1.0031 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | 3.1539e-07 | 0.00024642 | 0.9999 | 1.0007 | 0.00011928 | 0.9999 | -0.10055 | 0.12714 | -0.00045457 | 4.25e-07 | -0.10108 | 0.14502
| UUT secondary | 0.2001 | -1.0477e-07 | 3.151e-05 | 4.7508e-07 | 1.6937e-07 | 2.0766e-07 | 20.69 | 40.02 | 2.001 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -7.8494e-07 | 0.00023606 | 0.9999 | 1.0007 | 0.00011928 | 0.9999 | -0.10165 | 0.11679 | 0.00063915 | 0.00031669 | -0.10564 | 0.13379
| l - 1200At(5A) | 0.30058 | -1.0123e-06 | 4.59e-05 | 1.5476e-07 | 1.4135e-07 | 1.5279e-07 | 20.69 | 60.1164 | 3.0058 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -5.0486e-06 | 0.00022892 | 0.9999 | 1.0007 | 0.00011928 | 0.99989 | -0.10591 | 0.10964 | 0.0024746 | 0.00084746 | -0.10868 | 0.12709
| Std Range | 0.49938 | -1.5646e-06 | 7.2325e-05 | 6.3545e-06 | 8.4703e-22 | 4.3665e-07 | 20.69 | 99.8767 | 4.9938 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -4.6968e-06 | 0.00021711 | 0.9999 | 1.0007 | 0.00011928 | 0.99989 | -0.10556 | 0.097832 | 0.0082769 | 0.0025253 | -0.11268 | 0.11489
| 100/5 | 0.60106 | -7.1973e-08 | 8.506e-05 | 8.5057e-07 | 1.8464e-07 | 3.6448e-07 | 20.69 | 120.2127 | 6.0106 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -1.7951e-07 | 0.00021215 | 0.9999 | 1.0007 | 0.00011928 | 0.9999 | -0.10104 | 0.092867 | 0.012359 | 0.0037057 | -0.11156 | 0.11002
| 47 Hz | 0.024814 | 1.2838e-08 | 4.2134e-06 | 8.7969e-07 | 1.7809e-07 | 1.8531e-07 | 20.7 | 4.9628 | 0.24814 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | 7.7561e-07 | 0.00025454 | 0.9999 | 1.0007 | 7.7923e-05 | 0.9999 | -0.10009 | 0.17662 | -0.0007992 | -9.9232e-05
|  | 0.050249 | -2.3372e-08 | 8.3446e-06 | 2.8283e-08 | 8.412e-08 | 1.2812e-07 | 20.7 | 10.0498 | 0.50249 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -6.9726e-07 | 0.00024895 | 0.9999 | 1.0007 | 7.7923e-05 | 0.9999 | -0.10156 | 0.17103 | -0.00072955 | -7.9089e-05
|  | 0.10024 | -1.1168e-07 | 1.6103e-05 | 1.6087e-08 | 1.5556e-07 | 1.5798e-07 | 20.7 | 20.0481 | 1.0024 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -1.6701e-06 | 0.00024081 | 0.9999 | 1.0007 | 7.7923e-05 | 0.9999 | -0.10253 | 0.16289 | -0.00045507 | 2.8045e-07
|  | 0.20054 | -1.0883e-06 | 3.0599e-05 | 8.2902e-08 | 4.2648e-08 | 2.3871e-07 | 20.71 | 40.109 | 2.0054 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -8.1351e-06 | 0.00022873 | 0.9999 | 1.0007 | 7.7923e-05 | 0.99989 | -0.109 | 0.1508 | 0.00064566 | 0.00031858
|  | 0.30059 | -1.4718e-06 | 4.4606e-05 | 4.4265e-08 | 3.0137e-07 | 1.72e-07 | 20.71 | 60.1184 | 3.0059 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -7.3399e-06 | 0.00022246 | 0.9999 | 1.0007 | 7.7923e-05 | 0.99989 | -0.1082 | 0.14453 | 0.0024749 | 0.00084752
|  | 0.5006 | -2.4862e-06 | 7.0086e-05 | 2.6133e-07 | 4.2163e-07 | 5.3062e-07 | 20.71 | 100.1203 | 5.006 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -7.4451e-06 | 0.00020988 | 0.9999 | 1.0007 | 7.7923e-05 | 0.99989 | -0.10831 | 0.13196 | 0.0083213 | 0.0025381
|  | 0.60077 | -1.5646e-06 | 8.2195e-05 | 3.0798e-07 | 8.4703e-22 | 1.668e-07 | 20.71 | 120.1531 | 6.0077 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -3.9042e-06 | 0.0002051 | 0.9999 | 1.0007 | 7.7923e-05 | 0.9999 | -0.10477 | 0.12718 | 0.012346 | 0.003702
| 53 Hz | 0.30146 | -3.822e-07 | 4.4675e-05 | 2.6454e-08 | 2.9017e-07 | 2.815e-07 | 20.73 | 60.292 | 3.0146 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -1.9006e-06 | 0.00022216 | 0.9999 | 1.0007 | 7.7923e-05 | 0.9999 | -0.10276 | 0.14424 | Burden: | R Re | R Im
| 100 ohm burden test, 53 Hz | 0.30153 | -0.00026525 | 7.2837e-05 | 1.6365e-08 | 1.0245e-06 | 8.1315e-20 | 20.73 | 60.3067 | 3.0153 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -0.0013187 | 0.00036212 | 0.9999 | 1.0007 | 7.7923e-05 | 0.99858 | -1.4196 | 0.28419 |  | 0.13168 | -0.013995
|  | 0.30155 | -1.1553e-06 | 3.6111e-05 | 2.7975e-08 | 1.022e-07 | 4.1378e-08 | 20.74 | 60.3108 | 3.0155 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -5.743e-06 | 0.00017952 | 0.9999 | 1.0007 | 7.7923e-05 | 0.99989 | -0.1066 | 0.1016
| SR830 reads across 0.1 as above. (The 0.05 is in series on the high side) | [25 empty] | mean of 53 and 47 Hz
|  | 75/5 range, l - 1200At(5A), vs shunts and 500 AT on 100/5 range. | [6 empty] | add 0.05 ohms to burden | [15 empty] | heating | heating | heating corrected
| 2018-09-26 | [21 empty] | mA/A | mrad | 0.1 shunt | 0.2 shunt | mA/A | mrad
| 53 Hz | 0.024599 | 2.326e-09 | 4.449e-06 | 2.0767e-08 | 1.0115e-07 | 8.2634e-08 | 20.75 | 4.9198 | 0.24599 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | 1.4175e-07 | 0.00027113 | 0.9999 | 1.0007 | 0.00011928 | 0.9999 | -0.10072 | 0.15185 | -0.00079959 | -9.9344e-05 | -0.10078 | 0.17758
| UUT range | 0.050655 | -1.7465e-07 | 9.0643e-06 | 2.1345e-08 | 1.2297e-07 | 1.3544e-07 | 20.75 | 10.1311 | 0.50655 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -5.1685e-06 | 0.00026825 | 0.9999 | 1.0007 | 0.00011928 | 0.99989 | -0.10603 | 0.14897 | -0.00072805 | -7.8657e-05 | -0.10439 | 0.17048
| 75/5 | 0.10032 | -5.3102e-07 | 1.7338e-05 | 1.688e-08 | 1.6939e-07 | 6.2517e-08 | 20.75 | 20.0648 | 1.0032 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -7.9347e-06 | 0.00025907 | 0.9999 | 1.0007 | 0.00011928 | 0.99989 | -0.1088 | 0.13979 | -0.00045446 | 4.571e-07 | -0.10882 | 0.15984
| UUT secondary | 0.20044 | -1.2489e-06 | 3.2929e-05 | 1.1961e-08 | 1.9611e-07 | 6.5052e-08 | 20.75 | 40.0886 | 2.0044 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -9.3406e-06 | 0.00024627 | 0.9999 | 1.0007 | 0.00011928 | 0.99989 | -0.1102 | 0.12699 | 0.00064416 | 0.00031814 | -0.11029 | 0.1457
| l - 1200At(5A) | 0.3004 | -2.2116e-06 | 4.8189e-05 | 8.3608e-08 | 3.1644e-09 | 4.637e-08 | 20.75 | 60.0799 | 3.004 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -1.1037e-05 | 0.00024048 | 0.9999 | 1.0007 | 0.00011928 | 0.99989 | -0.1119 | 0.1212 | 0.0024706 | 0.0008463 | -0.11439 | 0.13901
| Std Range | 0.50078 | -3.2703e-06 | 7.6001e-05 | 5.5825e-07 | 2.7356e-07 | 9.5848e-08 | 20.75 | 100.1561 | 5.0078 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -9.7897e-06 | 0.00022751 | 0.9999 | 1.0007 | 0.00011928 | 0.99989 | -0.11065 | 0.10823 | 0.0083278 | 0.00254 | -0.11842 | 0.12664
| 100/5 | 0.60432 | -3.1293e-06 | 8.9867e-05 | 3.5986e-07 | 1.6941e-21 | 1.5798e-07 | 20.75 | 120.8646 | 6.0432 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -7.7625e-06 | 0.00022293 | 0.9999 | 1.0007 | 0.00011928 | 0.99989 | -0.10862 | 0.10365 | 0.012503 | 0.0037472 | -0.11891 | 0.12158
| 47 Hz | 0.024633 | -2.2757e-08 | 4.6212e-06 | 3.4598e-08 | 9.9716e-08 | 1.2309e-07 | 20.75 | 4.9266 | 0.24633 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -1.3849e-06 | 0.00028123 | 0.9999 | 1.0007 | 7.7923e-05 | 0.9999 | -0.10225 | 0.20331 | -0.00079953 | -9.9326e-05
|  | 0.050013 | -1.066e-07 | 9.0048e-06 | 3.0408e-08 | 1.0691e-07 | 1.0618e-07 | 20.75 | 10.0025 | 0.50013 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -3.1953e-06 | 0.00026991 | 0.9999 | 1.0007 | 7.7923e-05 | 0.9999 | -0.10406 | 0.19199 | -0.00073041 | -7.9339e-05
|  | 0.10053 | -5.9695e-07 | 1.729e-05 | 3.5422e-08 | 1.3941e-07 | 2.0021e-07 | 20.76 | 20.1061 | 1.0053 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -8.9017e-06 | 0.00025782 | 0.9999 | 1.0007 | 7.7923e-05 | 0.99989 | -0.10976 | 0.1799 | -0.00045294 | 8.9485e-07
|  | 0.2002 | -1.1852e-06 | 3.2364e-05 | 4.7685e-08 | 1.3595e-07 | 6.7946e-08 | 20.77 | 40.0404 | 2.002 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -8.8743e-06 | 0.00024234 | 0.9999 | 1.0007 | 7.7923e-05 | 0.99989 | -0.10974 | 0.16442 | 0.00064064 | 0.00031713
|  | 0.30037 | -2.5598e-06 | 4.7034e-05 | 2.2955e-08 | 2.3503e-07 | 1.1994e-07 | 20.77 | 60.0743 | 3.0037 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -1.2775e-05 | 0.00023474 | 0.9999 | 1.0007 | 7.7923e-05 | 0.99989 | -0.11364 | 0.15681 | 0.00247 | 0.00084613
|  | 0.5011 | -4.5979e-06 | 7.4531e-05 | 2.3203e-07 | 1.7958e-07 | 3.3757e-07 | 20.77 | 100.2203 | 5.011 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -1.3755e-05 | 0.00022297 | 0.9999 | 1.0007 | 7.7923e-05 | 0.99989 | -0.11462 | 0.14504 | 0.0083396 | 0.0025434
|  | 0.60078 | -4.3365e-06 | 8.7142e-05 | 3.0631e-07 | 3.4929e-07 | 1.912e-07 | 20.77 | 120.1556 | 6.0078 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -1.0821e-05 | 0.00021744 | 0.9999 | 1.0007 | 7.7923e-05 | 0.99989 | -0.11168 | 0.13952 | 0.012347 | 0.0037021
| 53 Hz | 0.30192 | -1.9536e-06 | 4.7388e-05 | 2.3273e-08 | 2.3824e-07 | 2.6738e-07 | 20.77 | 60.3844 | 3.0192 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -9.7e-06 | 0.00023529 | 0.9999 | 1.0007 | 7.7923e-05 | 0.99989 | -0.11056 | 0.15737 | Burden: | R Re | R Im
| 100 ohm burden test, 53 Hz | 0.30195 | -0.00036621 | 8.5852e-05 | 1.7619e-08 | 8.5653e-06 | 3.145e-06 | 20.78 | 60.3896 | 3.0195 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -0.0018181 | 0.00042623 | 0.9999 | 1.0007 | 7.7923e-05 | 0.99808 | -1.919 | 0.34831 |  | 0.18084 | -0.019094
|  | 0.30194 | -3.2795e-06 | 3.5154e-05 | 1.4436e-08 | 7.5206e-08 | 9.6083e-08 | 20.78 | 60.3882 | 3.0194 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -1.6282e-05 | 0.00017453 | 0.9999 | 1.0007 | 7.7923e-05 | 0.99988 | -0.11714 | 0.096612
| SR830 reads across 0.1 as above. (The 0.5 is in series on the high side) | [25 empty] | mean of 53 and 47 Hz
|  | 75/5 range, l - 1200At(5A), vs shunts and 500 AT on 100/5 range. | [6 empty] | add 0.5 ohms to burden (2 Tinsley 1659 1ohm in parallel) | [17 empty] | heating corrected
| 2018-09-26 | [25 empty] | mA/A | mrad
| 53 Hz | 0.024805 | -7.6241e-07 | 7.5094e-06 | 2.2886e-08 | 1.5265e-07 | 1.4845e-07 | 20.75 | 4.961 | 0.24805 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -4.6076e-05 | 0.00045383 | 0.9999 | 1.0007 | 0.00011928 | 0.99985 | -0.14694 | 0.33455 | -0.00079922 | -9.9236e-05 | -0.15087 | 0.36344
| UUT range | 0.050508 | -1.6699e-06 | 1.4748e-05 | 2.2995e-08 | 1.9045e-07 | 1.9568e-07 | 20.74 | 10.1016 | 0.50508 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -4.9564e-05 | 0.00043774 | 0.9999 | 1.0007 | 0.00011928 | 0.99985 | -0.15043 | 0.31846 | -0.00072859 | -7.8814e-05 | -0.15551 | 0.34142
| 75/5 | 0.10079 | -3.7761e-06 | 2.8132e-05 | 8.0493e-09 | 2.8164e-07 | 2.1183e-07 | 20.74 | 20.159 | 1.0079 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -5.616e-05 | 0.00041839 | 0.9999 | 1.0007 | 0.00011928 | 0.99984 | -0.15702 | 0.29911 | -0.000451 | 1.4564e-06 | -0.16148 | 0.3188
| UUT secondary | 0.20018 | -8.6369e-06 | 5.2456e-05 | 1.8342e-08 | 1.8431e-07 | 3.9005e-07 | 20.74 | 40.0363 | 2.0018 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -6.4679e-05 | 0.00039282 | 0.9999 | 1.0007 | 0.00011928 | 0.99983 | -0.16554 | 0.27355 | 0.00064034 | 0.00031704 | -0.17106 | 0.29116
| l - 1200At(5A) | 0.3007 | -1.3853e-05 | 7.4519e-05 | 5.2306e-08 | 1.7223e-08 | 3.8168e-08 | 20.75 | 60.1398 | 3.007 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -6.9062e-05 | 0.0003715 | 0.9999 | 1.0007 | 0.00011928 | 0.99983 | -0.16992 | 0.25222 | 0.0024772 | 0.0008482 | -0.17688 | 0.27084
| Std Range | 0.50021 | -2.4276e-05 | 0.00011305 | 1.8868e-07 | 1.4158e-08 | 8.9557e-07 | 20.75 | 100.0421 | 5.0021 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -7.2753e-05 | 0.00033879 | 0.9999 | 1.0007 | 0.00011928 | 0.99983 | -0.17361 | 0.21951 | 0.008307 | 0.002534 | -0.18589 | 0.2382
| 100/5 | 0.60093 | -3.0592e-05 | 0.00012844 | 3.1198e-07 | 5.3257e-08 | 1.0268e-06 | 20.75 | 120.185 | 6.0093 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -7.6316e-05 | 0.00032042 | 0.9999 | 1.0007 | 0.00011928 | 0.99982 | -0.17718 | 0.20114 | 0.012353 | 0.003704 | -0.18979 | 0.22037
| 47 Hz | 0.024752 | -9.1388e-07 | 7.7647e-06 | 2.6618e-08 | 1.1942e-07 | 1.6007e-07 | 20.77 | 4.9505 | 0.24752 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -5.5348e-05 | 0.00047025 | 0.9999 | 1.0007 | 7.7923e-05 | 0.99984 | -0.15621 | 0.39233
|  | 0.050445 | -2.0535e-06 | 1.4884e-05 | 2.6024e-08 | 7.6113e-08 | 4.9113e-08 | 20.77 | 10.0889 | 0.50445 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -6.1024e-05 | 0.00044231 | 0.9999 | 1.0007 | 7.7923e-05 | 0.99984 | -0.16189 | 0.36438
|  | 0.10074 | -4.434e-06 | 2.7982e-05 | 4.2069e-08 | 2.463e-08 | 1.1522e-07 | 20.77 | 20.1475 | 1.0074 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -6.5983e-05 | 0.00041641 | 0.9999 | 1.0007 | 7.7923e-05 | 0.99983 | -0.16684 | 0.33849
|  | 0.20017 | -1.0025e-05 | 5.1634e-05 | 3.2847e-08 | 6.6376e-08 | 1.803e-07 | 20.77 | 40.0337 | 2.0017 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -7.508e-05 | 0.00038669 | 0.9999 | 1.0007 | 7.7923e-05 | 0.99982 | -0.17594 | 0.30877
|  | 0.30078 | -1.5996e-05 | 7.3711e-05 | 3.9708e-08 | 8.099e-08 | 3.5727e-07 | 20.77 | 60.1565 | 3.0078 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -7.9725e-05 | 0.00036737 | 0.9999 | 1.0007 | 7.7923e-05 | 0.99982 | -0.18059 | 0.28945
|  | 0.50078 | -2.8646e-05 | 0.00011185 | 1.887e-07 | 1.0376e-07 | 3.3986e-07 | 20.77 | 100.157 | 5.0078 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -8.5753e-05 | 0.00033481 | 0.9999 | 1.0007 | 7.7923e-05 | 0.99981 | -0.18661 | 0.25688
|  | 0.6003 | -3.3732e-05 | 0.00012715 | 4.485e-07 | 3.738e-09 | 8.9872e-07 | 20.77 | 120.0599 | 6.003 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -8.4238e-05 | 0.00031753 | 0.9999 | 1.0007 | 7.7923e-05 | 0.99981 | -0.1851 | 0.2396
| 53 Hz | 0.30011 | -1.3874e-05 | 6.9657e-05 | 4.6989e-08 | 4.4286e-08 | 1.5553e-07 | 20.77 | 60.0225 | 3.0011 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -6.9301e-05 | 0.00034795 | 0.9999 | 1.0007 | 7.7923e-05 | 0.99983 | -0.17016 | 0.27002 | Burden: | R Re | R Im
| 100 ohm burden test, 53 Hz | 0.30046 | -0.001273 | 0.00018608 | 2.1794e-08 | 1.4028e-06 | 1.0193e-06 | 20.78 | 60.0914 | 3.0046 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -0.0063517 | 0.00092844 | 0.9999 | 1.0007 | 7.7923e-05 | 0.99355 | -6.4525 | 0.85051 |  | 0.62824 | -0.058049
|  | 0.30029 | -1.7572e-05 | 2.8063e-05 | 2.4137e-08 | 2.5132e-07 | 3.9578e-07 | 20.78 | 60.0583 | 3.0029 | 20 | 15 | 0.66707 | 0.2 | 0.1 | 1 | -8.7721e-05 | 0.0001401 | 0.9999 | 1.0007 | 7.7923e-05 | 0.99981 | -0.18858 | 0.062173